## OBSP Data
AnalyteName | Schedulecode | Parm+Method Code | Analysis Date | ReportUnit | RECOVERY | Login Date | Laboratory ID | ManufShortName | CALotNum | SpikeMix | Expected Conc. | newanalyte | Remark Code | REPORTING LEVEL | Lab Value | UNITS | All | Comments | Spike Level | Pct Recovery | Sample Spiked | STATUS | PLOT | perm_e | N | Mean | Std-Dev. | Median | IQR | F_Pseudo | Median Pct Recovery | maxout
FLUORANTHENE | 1433 | 34377GCM37 | 2012-08-07 | ug/L |  | 2012-06-14 06:00 | 20121670020 | Abs | 051112 | Abs 051112: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 79 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-01-29 | ug/L |  | 2009-01-15 08:30 | 20090160002 | Abs | 080508WW | Abs 080508WW: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-03-21 | ug/L |  | 2009-02-19 07:15 | 20090500129 | Abs | 080508WW | Abs 080508WW: 40 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-05-28 | ug/L |  | 2009-04-30 07:30 | 20091210001 | Abs | 080508WW | Abs 080508WW: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-08-06 | ug/L |  | 2009-07-02 09:00 | 20091830166 | Abs | 080508WW | Abs 080508WW: 40 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-10-09 | ug/L |  | 2009-08-20 08:00 | 20092330003 | Abs | 080508WW | Abs 080508WW: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-01-20 | ug/L |  | 2009-11-05 07:00 | 20093100004 | Accu | 209081217 | Accu 209081217: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-06-10 | ug/L |  | 2010-04-01 07:30 | 20100920049 | Accu | 209081217 | Accu 209081217: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-08-04 | ug/L |  | 2010-05-27 08:00 | 20101480009 | Accu | 209081217 | Accu 209081217: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-09-15 | ug/L |  | 2010-07-13 08:00 | 20101950139 | Accu | 209081217 | Accu 209081217: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-12-09 | ug/L |  | 2010-11-04 06:00 | 20103080117 | Accu | 210091144 | Accu 210091144: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-12-15 | ug/L |  | 2010-10-07 06:00 | 20102810021 | Accu | 210091144 | Accu 210091144: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-03-03 | ug/L |  | 2011-01-20 06:30 | 20110210097 | Accu | 210091144 | Accu 210091144: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-05-24 | ug/L |  | 2011-04-22 05:55 | 20111170110 | Accu | 210091144 | Accu 210091144: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-06-17 | ug/L |  | 2011-05-26 06:00 | 20111470063 | Accu | 210091144 | Accu 210091144: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-07-22 | ug/L |  | 2011-06-23 06:00 | 20111750005 | Accu | 210091144 | Accu 210091144: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-08-09 | ug/L |  | 2011-07-14 06:30 | 20111960040 | Accu | 210091144 | Accu 210091144: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-01-10 | ug/L |  | 2011-10-27 06:00 | 20113010044 | Accu | 210091144 | Accu 210091144: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-02-15 | ug/L |  | 2011-11-29 06:30 | 20113340045 | Accu | 210091144 | Accu 210091144: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-03-08 | ug/L |  | 2012-01-12 06:00 | 20120120070 | Accu | 210091144 | Accu 210091144: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-03-09 | ug/L |  | 2012-02-09 06:00 | 20120400057 | Accu | 210091144 | Accu 210091144: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-04-19 | ug/L |  | 2012-03-15 06:30 | 20120760114 | Accu | 210091144 | Accu 210091144: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-06-13 | ug/L |  | 2012-05-03 06:00 | 20121250042 | Accu | 210091144 | Accu 210091144: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-03-21 | ug/L |  | 2009-02-26 06:00 | 20090580005 | Crescent | 08HH017 | Crescent 08HH017: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-03-24 | ug/L |  | 2009-03-05 07:30 | 20090640122 | Crescent | 08HH017 | Crescent 08HH017: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-06-03 | ug/L |  | 2009-05-06 08:30 | 20091320006 | Crescent | 08HH017 | Crescent 08HH017: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-08-26 | ug/L |  | 2009-07-07 08:30 | 20091880137 | Crescent | 08HH017 | Crescent 08HH017: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-10-21 | ug/L |  | 2009-08-27 08:30 | 20092400049 | Crescent | 08HH017 | Crescent 08HH017: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-01-20 | ug/L |  | 2009-10-22 07:00 | 20092960004 | Crescent | 08LL076 | Crescent 08LL076: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-02-05 | ug/L |  | 2009-12-03 07:00 | 20093380006 | Crescent | 08LL076 | Crescent 08LL076: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-08-04 | ug/L |  | 2010-05-20 07:30 | 20101410008 | Crescent | 08LL076 | Crescent 08LL076: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-11-04 | ug/L |  | 2010-08-26 07:30 | 20102390089 | Crescent | 08LL076 | Crescent 08LL076: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-10-12 | ug/L |  | 2012-09-06 06:30 | 20122500215 | ERA | 0507-12-05.2 | ERA 0507-12-05.2: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-04-23 | ug/L |  | 2009-04-02 07:00 | 20090920065 | ERA | 0804-08-02WW | ERA 0804-08-02WW: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-07-02 | ug/L |  | 2009-05-28 05:50 | 20091490002 | ERA | 0804-08-02WW | ERA 0804-08-02WW: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-09-29 | ug/L |  | 2009-07-30 08:30 | 20092110207 | ERA | 0804-08-02WW | ERA 0804-08-02WW: 40 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-10-22 | ug/L |  | 2009-09-10 08:00 | 20092530079 | ERA | 0804-08-02WW | ERA 0804-08-02WW: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-01-20 | ug/L |  | 2009-11-02 06:30 | 20093060174 | Restek | A070024 | Restek A070024: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-02-23 | ug/L |  | 2009-12-10 06:30 | 20093450030 | Restek | A070024 | Restek A070024: 10 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-06-16 | ug/L |  | 2010-05-06 08:00 | 20101270034 | Restek | A070024 | Restek A070024: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-10-14 | ug/L |  | 2010-08-12 08:00 | 20102250101 | Restek | A070024 | Restek A070024: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-12-15 | ug/L |  | 2010-10-21 06:00 | 20102950074 | Restek | A075799 | Restek A075799: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-01-29 | ug/L |  | 2010-11-18 06:00 | 20103230119 | Restek | A075799 | Restek A075799: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-03-10 | ug/L |  | 2011-02-24 06:30 | 20110560003 | Restek | A075799 | Restek A075799: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-04-27 | ug/L |  | 2011-04-07 06:00 | 20110980032 | Restek | A075799 | Restek A075799: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-06-08 | ug/L |  | 2011-05-12 06:00 | 20111330147 | Restek | A075799 | Restek A075799: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-06-30 | ug/L |  | 2011-06-07 06:00 | 20111590003 | Restek | A075799 | Restek A075799: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-08-25 | ug/L |  | 2011-07-28 06:00 | 20112100006 | Restek | A075799 | Restek A075799: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-10-26 | ug/L |  | 2011-09-01 06:00 | 20112450025 | Restek | A075799 | Restek A075799: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-01-19 | ug/L |  | 2011-11-17 09:00 | 20113220200 | Restek | A075799 | Restek A075799: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-03-09 | ug/L |  | 2012-02-02 11:40 | 20120370004 | Restek | A075799 | Restek A075799: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-04-18 | ug/L |  | 2012-03-08 06:00 | 20120690029 | Restek | A075799 | Restek A075799: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-04-18 | ug/L |  | 2012-02-23 06:30 | 20120540034 | Restek | A075799 | Restek A075799: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-06-15 | ug/L |  | 2012-05-31 06:00 | 20121520249 | Restek | A075799 | Restek A075799: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-03-05 | ug/L |  | 2009-01-29 07:00 | 20090300002 | Supelco | LB60966 | Supelco LB60966: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-04-30 | ug/L |  | 2009-04-09 09:30 | 20090990158 | Supelco | LB60966 | Supelco LB60966: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-07-30 | ug/L |  | 2009-06-04 08:30 | 20091550183 | Supelco | LB60966 | Supelco LB60966: 40 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-09-30 | ug/L |  | 2009-08-06 08:00 | 20092180179 | Supelco | LB60966 | Supelco LB60966: 20 µl |  | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-12-01 | ug/L |  | 2009-09-17 08:00 | 20092610018 | Supelco | LB60966 | Supelco LB60966: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-01-15 | ug/L |  | 2009-10-15 06:30 | 20092890007 | Ultra | CF-3634 | Ultra CF-3634: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-01-27 | ug/L |  | 2009-11-19 06:30 | 20093240042 | Ultra | CF-3634 | Ultra CF-3634: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-03-25 | ug/L |  | 2010-03-11 07:00 | 20100710177 | Ultra | CF-3634 | Ultra CF-3634: 40 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-04-20 | ug/L |  | 2010-03-25 07:00 | 20100880031 | Ultra | CF-3634 | Ultra CF-3634: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-09-03 | ug/L |  | 2010-06-17 07:30 | 20101720148 | Ultra | CF-3634 | Ultra CF-3634: 80 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-12-05 | ug/L |  | 2012-10-18 06:00 | 20122930149 | Abs, Accu | 051112, 212051389 | Abs 051112: 100 µl & Accu 212051389: 100 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-08-30 | ug/L |  | 2012-07-26 06:00 | 20122090058 | Supelco, Accu | LB92784, 212051389 | Supelco LB92784: 10 µl & Accu 212051389: 10 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-01-23 | ug/L |  | 2012-11-29 06:30 | 20123350008 | Supelco, Accu | LB92784, 212051389 | Supelco LB92784: 10 µl & Accu 212051389: 10 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-04-16 | ug/L |  | 2013-03-14 06:30 | 20130740020 | Supelco, Accu | LB92784, 212051389 | Supelco LB92784: 50 µl & Accu 212051389: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-08-21 | ug/L |  | 2013-08-08 07:00 | 20132210040 | Abs, Accu | 040913WWF, 213041004 | Abs 040913WWF: 20 µl & Accu 213041004: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-05-09 | ug/L |  | 2014-03-06 08:30 | 20140650078 | Abs, Accu | 040913WWF, 213041004 | Abs 040913WWF: 100 µl & Accu 213041004: 100 µl |  | FLUORANTHENE | E | 0.024 | 0.002 | MICROGRAMS PER LITER |  |  |  |  | N | False Positives | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-12-05 | ug/L |  | 2013-10-17 06:00 | 20132910009 | Supelco, Accu | LB99396, 213041004 | Supelco LB99396: 20 µl & Accu 213041004: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-02-12 | ug/L |  | 2013-12-12 06:00 | 20133470006 | Abs, ERA | 040913WWF, 0318-13-01.3 | Abs 040913WWF: 20 µl & ERA 0318-13-01.3: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-07-11 | ug/L |  | 2013-06-20 07:00 | 20131720020 | Supelco, ERA | LB99396, 0318-13-01.3 | Supelco LB99396: 50 µl & ERA 0318-13-01.3: 50 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-05-09 | ug/L |  | 2014-02-06 06:00 | 20140370078 | Supelco, ERA | LB99396, 0318-13-01.3 | Supelco LB99396: 20 µl & ERA 0318-13-01.3: 20 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-02-28 | ug/L |  | 2013-01-31 06:30 | 20130320041 | Supelco, ERA | LB92784, 0507-12-05.2 | Supelco LB92784: 10 µl & ERA 0507-12-05.2: 10 µl |  | FLUORANTHENE | < | 0.024 | 0.024 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-01-29 | ug/L |  | 2009-01-21 08:00 | 20090210032 | Protocol | D1080815007 | Protocol D1080815007: 40 µl | 2.8161 | FLUORANTHENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | Y | False Negatives | N |  | 0 | [5 empty] | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-06-08 | ug/L | 1.0974 | 2011-05-19 06:30 | 20111390176 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 20 µl | 0.040096 | FLUORANTHENE | . | 0.024 | 0.044 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0974 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-12-09 | ug/L | 0.71995 | 2010-11-24 06:00 | 20103330009 | Ultra | CG-2301 | Ultra CG-2301: 20 µl | 0.077783 | FLUORANTHENE | . | 0.024 | 0.056 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.71995 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-06-19 | ug/L | 0.97468 | 2013-06-06 07:00 | 20131580039 | Crescent, Abs | 04CC145, 040913WWF | Crescent 04CC145: 20 µl & Abs 040913WWF: 20 µl | 0.079 | FLUORANTHENE | . | 0.024 | 0.077 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97468 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-08-20 | ug/L | 0.89873 | 2013-07-18 06:00 | 20132000020 | Crescent, Accu | 04CC145, 213041004 | Crescent 04CC145: 20 µl & Accu 213041004: 20 µl | 0.079 | FLUORANTHENE | . | 0.024 | 0.071 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.89873 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-05-09 | ug/L | 0.79747 | 2014-02-27 07:00 | 20140590010 | Crescent, Accu | 04CC145, 213041004 | Crescent 04CC145: 20 µl & Accu 213041004: 20 µl | 0.079 | FLUORANTHENE | . | 0.024 | 0.063 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79747 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-06-18 | ug/L | 1.0506 | 2013-05-23 06:30 | 20131440105 | Crescent, Supelco | 04CC145, LB99396 | Crescent 04CC145: 20 µl & Supelco LB99396: 20 µl | 0.079 | FLUORANTHENE | . | 0.024 | 0.083 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0506 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-11-08 | ug/L | 0.9625 | 2011-10-13 06:00 | 20112870076 | Ultra | CG-2301 | Ultra CG-2301: 20 µl | 0.08 | FLUORANTHENE | . | 0.024 | 0.077 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9625 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-04-11 | ug/L | 1.1875 | 2012-03-01 06:30 | 20120620014 | Ultra | CG-2301 | Ultra CG-2301: 20 µl | 0.08 | FLUORANTHENE | . | 0.024 | 0.095 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1875 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-02-12 | ug/L | 0.9875 | 2014-01-02 06:00 | 20140030012 | Spex Organics, Accu | T1130531001, 213041004 | Spex Organics T1130531001: 20 µl & Accu 213041004: 20 µl | 0.08 | FLUORANTHENE | . | 0.024 | 0.079 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9875 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-09-24 | ug/L | 0.95 | 2013-08-29 06:00 | 20132420002 | Spex Organics, ERA | T1130531001, 0318-13-01.3 | Spex Organics T1130531001: 20 µl & ERA 0318-13-01.3: 20 µl | 0.08 | FLUORANTHENE | . | 0.024 | 0.076 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-05-09 | ug/L | 0.975 | 2014-04-03 06:30 | 20140930051 | Spex Organics, ERA | T1130531001, 0318-13-01.3 | Spex Organics T1130531001: 20 µl & ERA 0318-13-01.3: 20 µl | 0.08 | FLUORANTHENE | . | 0.024 | 0.078 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.975 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-12-05 | ug/L | 0.8875 | 2013-09-19 06:00 | 20132630003 | Spex Organics, Supelco | T1130531001, LB99396 | Spex Organics T1130531001: 20 µl & Supelco LB99396: 20 µl | 0.08 | FLUORANTHENE | . | 0.024 | 0.071 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8875 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-05-09 | ug/L | 0.85 | 2014-03-13 06:30 | 20140730137 | Spex Organics, Supelco | T1130531001, LB99396 | Spex Organics T1130531001: 20 µl & Supelco LB99396: 20 µl | 0.08 | FLUORANTHENE | E | 0.024 | 0.068 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.85 | Y | Estimated Values | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-12-08 | ug/L | 0.96726 | 2010-10-14 06:30 | 20102880004 | Spex Organics | T1100706008 | Spex Organics T1100706008: 20 µl | 0.08064 | FLUORANTHENE | . | 0.024 | 0.078 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96726 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-06-29 | ug/L | 1.0137 | 2011-06-02 06:30 | 20111540003 | Ultra | CG-2301 | Ultra CG-2301: 20 µl | 0.08089 | FLUORANTHENE | . | 0.024 | 0.082 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0137 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-01-28 | ug/L | 1.137 | 2010-11-10 06:30 | 20103140134 | Spex Organics | T1100706008 | Spex Organics T1100706008: 20 µl | 0.080912 | FLUORANTHENE | . | 0.024 | 0.092 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.137 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-04-06 | ug/L | 1.0859 | 2011-03-17 06:30 | 20110770002 | Ultra | CG-2301 | Ultra CG-2301: 20 µl | 0.081037 | FLUORANTHENE | . | 0.024 | 0.088 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0859 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-08-25 | ug/L | 1.1765 | 2011-07-21 06:30 | 20112030054 | Spex Organics | T1100706008 | Spex Organics T1100706008: 20 µl | 0.0816 | FLUORANTHENE | . | 0.024 | 0.096 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1765 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-01-18 | ug/L | 1.0049 | 2011-11-03 06:30 | 20113080003 | Spex Organics | T1100706008 | Spex Organics T1100706008: 20 µl | 0.0816 | FLUORANTHENE | . | 0.024 | 0.082 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0049 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-03-09 | ug/L | 1.152 | 2012-01-26 06:00 | 20120260059 | Spex Organics | T1100706008 | Spex Organics T1100706008: 20 µl | 0.0816 | FLUORANTHENE | . | 0.024 | 0.094 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.152 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-08-08 | ug/L | 0.90909 | 2012-07-12 06:30 | 20121950065 | Crescent, Abs | 05BB047, 051112 | Crescent 05BB047: 10 µl & Abs 051112: 10 µl | 0.099 | FLUORANTHENE | . | 0.024 | 0.09 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90909 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-04-16 | ug/L | 0.94949 | 2013-02-28 06:30 | 20130590065 | Crescent, Abs | 05BB047, 051112 | Crescent 05BB047: 10 µl & Abs 051112: 10 µl | 0.099 | FLUORANTHENE | . | 0.024 | 0.094 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94949 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-01-23 | ug/L | 0.78788 | 2012-12-13 06:30 | 20123490152 | Crescent, ERA | 05BB047, 0507-12-05.2 | Crescent 05BB047: 10 µl & ERA 0507-12-05.2: 10 µl | 0.099 | FLUORANTHENE | . | 0.024 | 0.078 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78788 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-08-26 | ug/L | 1.2 | 2011-08-10 06:30 | 20112220131 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 50 µl | 0.1 | FLUORANTHENE | . | 0.024 | 0.12 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-11-08 | ug/L | 1.1 | 2011-10-07 07:00 | 20112840029 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 50 µl | 0.1 | FLUORANTHENE | . | 0.024 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-12-05 | ug/L | 0.87 | 2012-11-15 06:30 | 20123200135 | Spex Organics, Abs | T1120521009, 051112 | Spex Organics T1120521009: 10 µl & Abs 051112: 10 µl | 0.1 | FLUORANTHENE | . | 0.024 | 0.087 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.87 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-08-30 | ug/L | 0.96 | 2012-08-09 06:30 | 20122230022 | Spex Organics, ERA | T1120521009, 0507-12-05.2 | Spex Organics T1120521009: 10 µl & ERA 0507-12-05.2: 10 µl | 0.1 | FLUORANTHENE | . | 0.024 | 0.096 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-08-31 | ug/L | 0.69644 | 2010-06-24 10:00 | 20101760004 | NWQL | BNA2SP094D | NWQL BNA2SP094D: 60 µl | 0.11918 | FLUORANTHENE | E | 0.024 | 0.083 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.69644 | Y | Estimated Values | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-06-04 | ug/L | 0.91813 | 2010-04-22 09:30 | 20101130052 | NWQL | BNA2SP094D | NWQL BNA2SP094D: 60 µl | 0.11981 | FLUORANTHENE | . | 0.024 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91813 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-08-02 | ug/L | 0.9375 | 2011-06-29 06:30 | 20111810002 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 80 µl | 0.16 | FLUORANTHENE | . | 0.024 | 0.15 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9375 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-12-05 | ug/L | 0.96203 | 2013-10-31 06:00 | 20133050051 | Crescent, ERA | 04CC145, 0318-13-01.3 | Crescent 04CC145: 50 µl & ERA 0318-13-01.3: 50 µl | 0.1975 | FLUORANTHENE | . | 0.024 | 0.19 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96203 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-08-26 | ug/L | 1 | 2011-08-04 06:30 | 20112170171 | Ultra | CG-2301 | Ultra CG-2301: 50 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-10-27 | ug/L | 0.85 | 2011-09-08 06:30 | 20112520206 | Ultra | CG-2301 | Ultra CG-2301: 50 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.17 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.85 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-02-08 | ug/L | 1.05 | 2011-11-22 06:00 | 20113260109 | Ultra | CG-2301 | Ultra CG-2301: 50 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.21 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.05 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-12-05 | ug/L | 1 | 2013-11-14 07:00 | 20133190011 | Spex Organics, Abs | T1130531001, 040913WWF | Spex Organics T1130531001: 50 µl & Abs 040913WWF: 50 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-10-11 | ug/L | 1 | 2012-08-23 06:00 | 20122370104 | Spex Organics, Accu | T1120521009, 212051389 | Spex Organics T1120521009: 20 µl & Accu 212051389: 20 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-05-15 | ug/L | 0.95 | 2013-04-11 06:00 | 20131020083 | Abs, Spex Organics | 051112, T1120521009 | Abs 051112: 20 µl & Spex Organics T1120521009: 20 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.19 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-08-08 | ug/L | 1.1 | 2012-06-28 06:30 | 20121810018 | Spex Organics, Supelco | T1120521009, LB92784 | Spex Organics T1120521009: 20 µl & Supelco LB92784: 20 µl | 0.2 | FLUORANTHENE | . | 0.024 | 0.22 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-06-08 | ug/L | 0.9461 | 2011-05-05 06:00 | 20111260155 | Ultra | CG-2301 | Ultra CG-2301: 50 µl | 0.20082 | FLUORANTHENE | . | 0.024 | 0.19 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9461 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-03-09 | ug/L | 0.99265 | 2011-02-10 06:30 | 20110410062 | Spex Organics | T1100706008 | Spex Organics T1100706008: 50 µl | 0.20148 | FLUORANTHENE | . | 0.024 | 0.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99265 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-11-04 | ug/L | 0.93137 | 2011-09-15 06:00 | 20112590003 | Spex Organics | T1100706008 | Spex Organics T1100706008: 50 µl | 0.204 | FLUORANTHENE | . | 0.024 | 0.19 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93137 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-02-16 | ug/L | 1.0294 | 2011-12-13 06:00 | 20113490055 | Spex Organics | T1100706008 | Spex Organics T1100706008: 50 µl | 0.204 | FLUORANTHENE | . | 0.024 | 0.21 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0294 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-05-08 | ug/L | 0.88235 | 2012-03-29 06:30 | 20120930006 | Spex Organics | T1100706008 | Spex Organics T1100706008: 50 µl | 0.204 | FLUORANTHENE | . | 0.024 | 0.18 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88235 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-05-09 | ug/L | 0.83333 | 2012-04-19 06:00 | 20121110021 | Spex Organics | T1100706008 | Spex Organics T1100706008: 50 µl | 0.204 | FLUORANTHENE | . | 0.024 | 0.17 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83333 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-12-16 | ug/L | 0.82867 | 2010-10-28 06:30 | 20103020151 | Ultra | CG-2301 | Ultra CG-2301: 80 µl | 0.31376 | FLUORANTHENE | . | 0.024 | 0.26 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82867 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-10-26 | ug/L | 0.84375 | 2011-08-25 07:00 | 20112380051 | Ultra | CG-2301 | Ultra CG-2301: 80 µl | 0.32 | FLUORANTHENE | . | 0.024 | 0.27 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84375 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-02-16 | ug/L | 0.90625 | 2011-12-15 06:30 | 20113500020 | Ultra | CG-2301 | Ultra CG-2301: 80 µl | 0.32 | FLUORANTHENE | . | 0.024 | 0.29 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90625 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-06-14 | ug/L | 0.875 | 2012-05-17 06:30 | 20121390008 | Ultra | CG-2301 | Ultra CG-2301: 80 µl | 0.32 | FLUORANTHENE | . | 0.024 | 0.28 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.875 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-04-28 | ug/L | 1.0137 | 2011-04-14 06:00 | 20111050152 | Spex Organics | T1100706008 | Spex Organics T1100706008: 80 µl | 0.32555 | FLUORANTHENE | . | 0.024 | 0.33 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0137 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2011-07-22 | ug/L | 0.8569 | 2011-06-16 06:30 | 20111680032 | Spex Organics | T1100706008 | Spex Organics T1100706008: 80 µl | 0.32676 | FLUORANTHENE | . | 0.024 | 0.28 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8569 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-02-03 | ug/L | 0.79484 | 2009-11-25 06:30 | 20093340133 | NSI | 083109-01WW | NSI 083109-01WW: 20 µl | 0.39002 | FLUORANTHENE | . | 0.024 | 0.31 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79484 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2014-02-12 | ug/L | 0.86076 | 2014-01-16 07:00 | 20140160087 | Crescent, Abs | 04CC145, 040913WWF | Crescent 04CC145: 100 µl & Abs 040913WWF: 100 µl | 0.395 | FLUORANTHENE | . | 0.024 | 0.34 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86076 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-10-14 | ug/L | 0.77979 | 2010-08-03 08:00 | 20102160003 | ChemService | 480-95B | ChemService 480-95B: 20 µl | 0.39754 | FLUORANTHENE | . | 0.024 | 0.31 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77979 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-06-03 | ug/L | 0.79433 | 2010-04-15 07:00 | 20101060001 | ChemService | 480-95B | ChemService 480-95B: 20 µl | 0.40285 | FLUORANTHENE | . | 0.024 | 0.32 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79433 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-11-04 | ug/L | 0.84328 | 2009-10-08 07:00 | 20092820006 | NSI | 083109-01WW | NSI 083109-01WW: 20 µl | 0.40319 | FLUORANTHENE | . | 0.024 | 0.34 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84328 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-02-27 | ug/L | 0.72727 | 2013-01-17 06:00 | 20130180028 | Crescent, Abs | 05BB047, 051112 | Crescent 05BB047: 50 µl & Abs 051112: 50 µl | 0.495 | FLUORANTHENE | . | 0.024 | 0.36 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72727 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-10-11 | ug/L | 0.80808 | 2012-09-19 06:00 | 20122640133 | Crescent, Accu | 05BB047, 212051389 | Crescent 05BB047: 50 µl & Accu 212051389: 50 µl | 0.495 | FLUORANTHENE | . | 0.024 | 0.4 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80808 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-05-15 | ug/L | 0.84848 | 2013-04-25 06:00 | 20131160078 | Crescent, ERA | 05BB047, 0507-12-05.2 | Crescent 05BB047: 50 µl & ERA 0507-12-05.2: 50 µl | 0.495 | FLUORANTHENE | . | 0.024 | 0.42 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84848 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2013-02-28 | ug/L | 0.7 | 2013-02-14 06:30 | 20130460004 | Spex Organics, ERA | T1120521009, 0507-12-05.2 | Spex Organics T1120521009: 50 µl & ERA 0507-12-05.2: 50 µl | 0.5 | FLUORANTHENE | . | 0.024 | 0.35 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-07-30 | ug/L | 0.78412 | 2009-07-16 08:30 | 20091970071 | NWQL | BNA2SP053B | NWQL BNA2SP053B: 75 µl | 0.5994 | FLUORANTHENE | . | 0.04 | 0.47 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78412 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-05-06 | ug/L | 0.8768 | 2009-04-16 08:30 | 20091060130 | NWQL | BNA2SP053B | NWQL BNA2SP053B: 75 µl | 0.60447 | FLUORANTHENE | . | 0.04 | 0.53 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8768 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-03-06 | ug/L | 0.73976 | 2009-02-05 09:30 | 20090370006 | NWQL | BNA2SP071C | NWQL BNA2SP071C: 75 µl | 0.62183 | FLUORANTHENE | . | 0.04 | 0.46 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.73976 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-04-02 | ug/L | 0.85968 | 2009-03-12 08:30 | 20090710011 | NWQL | BNA2SP053B | NWQL BNA2SP053B: 75 µl | 0.62814 | FLUORANTHENE | . | 0.04 | 0.54 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.85968 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-01-21 | ug/L | 0.78203 | 2009-11-12 09:00 | 20093170003 | ChemService | 480-95B | ChemService 480-95B: 40 µl | 0.7928 | FLUORANTHENE | . | 0.024 | 0.62 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78203 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-03-25 | ug/L | 0.74369 | 2010-02-18 07:00 | 20100500004 | NSI | 083109-01WW | NSI 083109-01WW: 40 µl | 0.79334 | FLUORANTHENE | . | 0.024 | 0.59 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.74369 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-06-29 | ug/L | 0.62245 | 2010-06-03 08:00 | 20101550005 | NSI | 083109-01WW | NSI 083109-01WW: 40 µl | 0.81934 | FLUORANTHENE | . | 0.024 | 0.51 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.62246 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-12-06 | ug/L | 0.9596 | 2012-10-04 06:00 | 20122790025 | Crescent, Supelco | 05BB047, LB92784 | Crescent 05BB047: 100 µl & Supelco LB92784: 100 µl | 0.99 | FLUORANTHENE | . | 0.024 | 0.95 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9596 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2012-12-07 | ug/L | 0.97 | 2012-11-01 06:00 | 20123070010 | Spex Organics, ERA | T1120521009, 0507-12-05.2 | Spex Organics T1120521009: 100 µl & ERA 0507-12-05.2: 100 µl | 1 | FLUORANTHENE | . | 0.024 | 0.97 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-08-27 | ug/L | 0.70424 | 2009-07-23 08:30 | 20092040204 | Protocol | D1080815007 | Protocol D1080815007: 20 µl | 1.3348 | FLUORANTHENE | . | 0.04 | 0.94 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70424 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-07-22 | ug/L | 0.82012 | 2009-05-20 08:30 | 20091410002 | Protocol | D1080815007 | Protocol D1080815007: 20 µl | 1.3413 | FLUORANTHENE | . | 0.04 | 1.1 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82012 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-10-07 | ug/L | 0.74234 | 2009-08-13 08:30 | 20092260011 | ChemService | 406-104C | ChemService 406-104C: 20 µl | 1.3471 | FLUORANTHENE | . | 0.024 | 1 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.74234 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-08-05 | ug/L | 0.81177 | 2009-06-18 09:00 | 20091700146 | ChemService | 406-104C | ChemService 406-104C: 20 µl | 1.3551 | FLUORANTHENE | . | 0.04 | 1.1 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81177 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-04-02 | ug/L | 0.95385 | 2009-03-19 08:30 | 20090780143 | Protocol | D1080815007 | Protocol D1080815007: 20 µl | 1.3629 | FLUORANTHENE | . | 0.04 | 1.3 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95385 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-03-05 | ug/L | 0.83138 | 2009-02-12 07:30 | 20090430170 | ChemService | 406-104C | ChemService 406-104C: 20 µl | 1.4073 | FLUORANTHENE | . | 0.04 | 1.17 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83138 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-01-27 | ug/L | 0.7748 | 2009-01-08 08:00 | 20090080049 | ChemService | 406-104C | ChemService 406-104C: 20 µl | 1.4197 | FLUORANTHENE | . | 0.04 | 1.1 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7748 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-03-25 | ug/L | 0.75675 | 2010-02-11 07:00 | 20100420054 | NSI | 083109-01WW | NSI 083109-01WW: 80 µl | 1.5857 | FLUORANTHENE | . | 0.024 | 1.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75675 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2010-02-23 | ug/L | 0.80026 | 2009-12-17 08:00 | 20093520065 | ChemService | 480-95B | ChemService 480-95B: 80 µl | 1.6245 | FLUORANTHENE | . | 0.024 | 1.3 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80026 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-10-22 | ug/L | 0.81182 | 2009-09-03 07:30 | 20092460252 | Protocol | D1080815007 | Protocol D1080815007: 40 µl | 2.71 | FLUORANTHENE | . | 0.024 | 2.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81182 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875
FLUORANTHENE | 1433 | 34377GCM37 | 2009-05-12 | ug/L | 0.84254 | 2009-04-23 08:30 | 20091130128 | ChemService | 406-104C | ChemService 406-104C: 40 µl | 2.7298 | FLUORANTHENE | . | 0.04 | 2.3 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84254 | Y | Recoverable Result | Y |  | 79 | 0.90381 | 0.12678 | 0.8875 | 0.17942 | 0.133 | 0.8875