## OBSP Data
AnalyteName | Schedulecode | Parm+Method Code | Analysis Date | ReportUnit | RECOVERY | Login Date | Laboratory ID | ManufShortName | CALotNum | SpikeMix | Expected Conc. | newanalyte | Remark Code | REPORTING LEVEL | Lab Value | UNITS | All | Comments | Spike Level | Pct Recovery | Sample Spiked | STATUS | PLOT | perm_e | N | Mean | Std-Dev. | Median | IQR | F_Pseudo | Median Pct Recovery | maxout
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-06-12 | ug/L |  | 2012-06-07 07:30 | 20121600154 | Abs | 051712 | Abs 051712: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 69 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-11-21 | ug/L |  | 2012-11-15 07:30 | 20123200133 | Abs | 051712 | Abs 051712: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-01-27 | ug/L |  | 2009-01-21 07:30 | 20090210031 | Abs | 080408VOC | Abs 080408VOC: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-03-14 | ug/L |  | 2009-03-05 07:00 | 20090640127 | Abs | 080408VOC | Abs 080408VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-04-21 | ug/L |  | 2009-04-16 08:00 | 20091060129 | Abs | 080408VOC | Abs 080408VOC: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-06-09 | ug/L |  | 2009-06-04 08:00 | 20091550189 | Abs | 080408VOC | Abs 080408VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-07-17 | ug/L |  | 2009-07-16 08:00 | 20091970076 | Abs | 080408VOC | Abs 080408VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-08-29 | ug/L |  | 2009-08-27 08:00 | 20092400054 | Abs | 080408VOC | Abs 080408VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-11-14 | ug/L |  | 2009-11-12 07:00 | 20093170002 | Abs | 090809 | Abs 090809: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-01-26 | ug/L |  | 2010-01-07 06:00 | 20100110075 | Abs | 090809 | Abs 090809: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-02-19 | ug/L |  | 2010-02-18 06:30 | 20100500006 | Abs | 090809 | Abs 090809: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-04-02 | ug/L |  | 2010-03-25 06:00 | 20100880030 | Abs | 090809 | Abs 090809: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-05-11 | ug/L |  | 2010-05-06 06:00 | 20101270040 | Abs | 090809 | Abs 090809: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-06-24 | ug/L |  | 2010-06-17 06:00 | 20101720156 | Abs | 090809 | Abs 090809: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-07-27 | ug/L |  | 2010-07-21 06:00 | 20102030002 | Abs | 090809 | Abs 090809: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-09-08 | ug/L |  | 2010-09-02 06:00 | 20102500044 | Abs | 090809 | Abs 090809: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-11-10 | ug/L |  | 2009-11-05 08:00 | 20093100001 | Accu | 209081198 | Accu 209081198: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-12-24 | ug/L |  | 2009-12-17 06:30 | 20093560038 | Accu | 209081198 | Accu 209081198: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-02-19 | ug/L |  | 2010-02-11 06:00 | 20100420058 | Accu | 209081198 | Accu 209081198: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-03-26 | ug/L |  | 2010-03-19 06:00 | 20100810066 | Accu | 209081198 | Accu 209081198: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-05-05 | ug/L |  | 2010-04-29 06:00 | 20101190085 | Accu | 209081198 | Accu 209081198: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-06-17 | ug/L |  | 2010-06-10 06:00 | 20101650035 | Accu | 209081198 | Accu 209081198: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-07-16 | ug/L |  | 2010-07-13 06:00 | 20101950135 | Accu | 209081198 | Accu 209081198: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-09-01 | ug/L |  | 2010-08-26 06:00 | 20102390264 | Accu | 209081198 | Accu 209081198: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-10-23 | ug/L |  | 2010-10-21 07:00 | 20102950069 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-12-03 | ug/L |  | 2010-11-30 06:30 | 20103350005 | Crescent | 06MM148 | Crescent 06MM148: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-02-14 | ug/L |  | 2011-02-02 06:00 | 20110340011 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-03-26 | ug/L |  | 2011-03-24 06:00 | 20110840074 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-05-10 | ug/L |  | 2011-05-05 07:00 | 20111260150 | Crescent | 06MM148 | Crescent 06MM148: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-06-18 | ug/L |  | 2011-06-16 07:30 | 20111680038 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-08-10 | ug/L |  | 2011-08-04 07:30 | 20112170149 | Crescent | 06MM148 | Crescent 06MM148: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-10-24 | ug/L |  | 2011-10-20 06:30 | 20112930089 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-12-17 | ug/L |  | 2011-12-13 07:00 | 20113490067 | Crescent | 06MM148 | Crescent 06MM148: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-02-24 | ug/L |  | 2012-02-16 07:00 | 20120470076 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-03-24 | ug/L |  | 2012-03-15 07:30 | 20120760109 | Crescent | 06MM148 | Crescent 06MM148: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-02-06 | ug/L |  | 2009-01-29 06:30 | 20090300007 | ERA | 0804-08-02VOC | ERA 0804-08-02VOC: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-03-14 | ug/L |  | 2009-03-12 08:00 | 20090710018 | ERA | 0804-08-02VOC | ERA 0804-08-02VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-04-25 | ug/L |  | 2009-04-23 08:00 | 20091130126 | ERA | 0804-08-02VOC | ERA 0804-08-02VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-06-17 | ug/L |  | 2009-06-11 09:00 | 20091620169 | ERA | 0804-08-02VOC | ERA 0804-08-02VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-07-25 | ug/L |  | 2009-07-23 08:00 | 20092040199 | ERA | 0804-08-02VOC | ERA 0804-08-02VOC: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-09-05 | ug/L |  | 2009-09-03 07:00 | 20092460254 | ERA | 0804-08-02VOC | ERA 0804-08-02VOC: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-11-05 | ug/L |  | 2010-11-04 07:00 | 20103080123 | NSI | 062310-01 | NSI 062310-01: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-01-22 | ug/L |  | 2011-01-13 06:00 | 20110130057 | NSI | 062310-01 | NSI 062310-01: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-03-01 | ug/L |  | 2011-02-24 07:30 | 20110560001 | NSI | 062310-01 | NSI 062310-01: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-04-13 | ug/L |  | 2011-04-07 06:30 | 20110980041 | NSI | 062310-01 | NSI 062310-01: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-05-03 | ug/L |  | 2011-04-28 06:30 | 20111190185 | NSI | 062310-01 | NSI 062310-01: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-06-10 | ug/L |  | 2011-06-07 07:00 | 20111590001 | NSI | 062310-01 | NSI 062310-01: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-09-17 | ug/L |  | 2011-09-15 07:00 | 20112590007 | NSI | 062310-01 | NSI 062310-01: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-10-18 | ug/L |  | 2011-10-13 07:00 | 20112870072 | NSI | 062310-01 | NSI 062310-01: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-11-24 | ug/L |  | 2011-11-22 07:00 | 20113260106 | NSI | 062310-01 | NSI 062310-01: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-03-10 | ug/L |  | 2012-03-08 06:30 | 20120690027 | NSI | 062310-01 | NSI 062310-01: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-05-08 | ug/L |  | 2012-05-03 06:30 | 20121250044 | NSI | 062310-01 | NSI 062310-01: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-05-25 | ug/L |  | 2012-05-24 06:30 | 20121460003 | NSI | 062310-01 | NSI 062310-01: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-10-24 | ug/L |  | 2009-10-22 08:00 | 20092960001 | NSI | 090909-01VOC | NSI 090909-01VOC: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-12-11 | ug/L |  | 2009-12-03 08:00 | 20093380012 | NSI | 090909-01VOC | NSI 090909-01VOC: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-02-03 | ug/L |  | 2010-01-28 06:00 | 20100280080 | NSI | 090909-01VOC | NSI 090909-01VOC: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-03-11 | ug/L |  | 2010-03-03 06:00 | 20100630040 | NSI | 090909-01VOC | NSI 090909-01VOC: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-04-21 | ug/L |  | 2010-04-15 06:00 | 20101060009 | NSI | 090909-01VOC | NSI 090909-01VOC: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-05-29 | ug/L |  | 2010-05-27 06:00 | 20101480004 | NSI | 090909-01VOC | NSI 090909-01VOC: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-07-09 | ug/L |  | 2010-07-08 06:30 | 20101900010 | NSI | 090909-01VOC | NSI 090909-01VOC: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-08-14 | ug/L |  | 2010-08-12 06:00 | 20102250095 | NSI | 090909-01VOC | NSI 090909-01VOC: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-10-22 | ug/L |  | 2009-10-15 08:00 | 20092890012 | Restek | A069860 | Restek A069860: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-12-04 | ug/L |  | 2009-11-25 07:30 | 20093340136 | Restek | A069860 | Restek A069860: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-01-26 | ug/L |  | 2010-01-21 06:00 | 20100210101 | Restek | A069860 | Restek A069860: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-03-11 | ug/L |  | 2010-02-25 06:00 | 20100570085 | Restek | A069860 | Restek A069860: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-04-13 | ug/L |  | 2010-04-08 06:00 | 20100990027 | Restek | A069860 | Restek A069860: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-05-22 | ug/L |  | 2010-05-20 06:00 | 20101410003 | Restek | A069860 | Restek A069860: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-07-09 | ug/L |  | 2010-07-01 06:00 | 20101830136 | Restek | A069860 | Restek A069860: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-08-11 | ug/L |  | 2010-08-05 06:00 | 20102180086 | Restek | A069860 | Restek A069860: 30 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-09-21 | ug/L |  | 2010-09-16 06:15 | 20102630175 | Restek | A069860 | Restek A069860: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-11-16 | ug/L |  | 2010-11-10 07:00 | 20103140133 | Restek | A075493 | Restek A075493: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-01-22 | ug/L |  | 2011-01-20 07:00 | 20110210094 | Restek | A075493 | Restek A075493: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-03-08 | ug/L |  | 2011-03-03 06:30 | 20110620085 | Restek | A075493 | Restek A075493: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-04-21 | ug/L |  | 2011-04-14 07:00 | 20111050149 | Restek | A075493 | Restek A075493: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-05-28 | ug/L |  | 2011-05-26 07:00 | 20111470062 | Restek | A075493 | Restek A075493: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-07-20 | ug/L |  | 2011-07-14 07:30 | 20111950194 | Restek | A075493 | Restek A075493: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-09-03 | ug/L |  | 2011-09-01 07:00 | 20112450030 | Restek | A075493 | Restek A075493: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-11-18 | ug/L |  | 2011-11-09 06:30 | 20113130089 | Restek | A075493 | Restek A075493: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-01-26 | ug/L |  | 2012-01-19 06:30 | 20120200074 | Restek | A075493 | Restek A075493: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-02-24 | ug/L |  | 2012-02-23 07:00 | 20120540039 | Restek | A075493 | Restek A075493: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-04-06 | ug/L |  | 2012-03-29 07:30 | 20120930001 | Restek | A075493 | Restek A075493: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-02-27 | ug/L |  | 2009-02-19 07:00 | 20090500156 | Supelco | LB60727 | Supelco LB60727: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-04-08 | ug/L |  | 2009-04-02 06:45 | 20090920063 | Supelco | LB60727 | Supelco LB60727: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-05-16 | ug/L |  | 2009-05-14 09:30 | 20091340076 | Supelco | LB60727 | Supelco LB60727: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-07-08 | ug/L |  | 2009-07-02 08:30 | 20091830175 | Supelco | LB60727 | Supelco LB60727: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-08-19 | ug/L |  | 2009-08-13 08:00 | 20092260015 | Supelco | LB60727 | Supelco LB60727: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-09-26 | ug/L |  | 2009-09-17 07:30 | 20092610022 | Supelco | LB60727 | Supelco LB60727: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-01-17 | ug/L |  | 2009-01-08 07:30 | 20090080050 | Ultra | CE-0471 | Ultra CE-0471: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-02-18 | ug/L |  | 2009-02-12 07:00 | 20090430174 | Ultra | CE-2869 | Ultra CE-2869: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-03-31 | ug/L |  | 2009-03-26 08:00 | 20090860070 | Ultra | CE-2869 | Ultra CE-2869: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-05-13 | ug/L |  | 2009-05-06 08:00 | 20091320008 | Ultra | CE-2869 | Ultra CE-2869: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-06-17 | ug/L |  | 2009-06-11 08:30 | 20091620168 | Ultra | CE-2869 | Ultra CE-2869: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-06-30 | ug/L |  | 2009-06-25 08:30 | 20091770124 | Ultra | CE-2869 | Ultra CE-2869: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-08-11 | ug/L |  | 2009-08-06 07:30 | 20092180176 | Ultra | CE-2869 | Ultra CE-2869: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-06-08 | ug/L |  | 2013-06-06 07:30 | 20131580021 | Accu, Abs | 213041120, 051413 | Accu 213041120: 10 µl & Abs 051413: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2014-03-19 | ug/L |  | 2014-03-06 07:30 | 20140650083 | Accu, Abs | 213041120, 051413 | Accu 213041120: 10 µl & Abs 051413: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-08-06 | ug/L |  | 2013-08-01 07:30 | 20132140003 | Supelco, Abs | LC00154, 051413 | Supelco LC00154: 20 µl & Abs 051413: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-12-12 | ug/L |  | 2013-12-05 07:30 | 20133400006 | Supelco, Abs | LC00154, 051413 | Supelco LC00154: 10 µl & Abs 051413: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-01-19 | ug/L |  | 2013-01-17 07:00 | 20130180025 | Accu, Abs | 212051193, 051712 | Accu 212051193: 40 µl & Abs 051712: 40 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-07-30 | ug/L |  | 2012-07-26 07:00 | 20122090060 | Supelco, Abs | LB92787, 051712 | Supelco LB92787: 20 µl & Abs 051712: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-11-01 | ug/L |  | 2013-10-24 07:30 | 20132980025 | Accu, ChemService | 213041120, 1530700 | Accu 213041120: 20 µl & ChemService 1530700: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-12-31 | ug/L |  | 2013-12-23 07:30 | 20133580039 | Accu, ChemService | 213041120, 1530700 | Accu 213041120: 10 µl & ChemService 1530700: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-09-17 | ug/L |  | 2013-09-12 07:30 | 20132560004 | Supelco, ChemService | LC00154, 1530700 | Supelco LC00154: 20 µl & ChemService 1530700: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2014-01-29 | ug/L |  | 2014-01-23 07:00 | 20140230083 | Supelco, ChemService | LC00154, 1530700 | Supelco LC00154: 20 µl & ChemService 1530700: 20 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-03-29 | ug/L |  | 2013-03-21 06:00 | 20130810110 | Supelco, ChemService | LB92787, 472200 | Supelco LB92787: 10 µl & ChemService 472200: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-05-03 | ug/L |  | 2013-04-25 06:30 | 20131160076 | Supelco, ChemService | LB92787, 472200 | Supelco LB92787: 10 µl & ChemService 472200: 10 µl |  | ETHYL_TERT-BUTYL_ETHER | < | 0.032 | 0.032 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-01-08 | ug/L | 0.7625 | 2011-01-06 06:30 | 20110070107 | Accu | 210061400 | Accu 210061400: 20 µl | 0.04 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0305 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7625 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-02-19 | ug/L | 0.98 | 2011-02-17 06:00 | 20110490036 | Accu | 210061400 | Accu 210061400: 20 µl | 0.04 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0392 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.98 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-04-29 | ug/L | 1.005 | 2011-04-22 06:15 | 20111170106 | Accu | 210061400 | Accu 210061400: 20 µl | 0.04 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0402 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.005 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-07-23 | ug/L | 1.06 | 2011-07-21 07:00 | 20112030060 | Accu | 210061400 | Accu 210061400: 20 µl | 0.04 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0424 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.06 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-11-03 | ug/L | 0.905 | 2011-10-27 07:00 | 20113010043 | Accu | 210061400 | Accu 210061400: 20 µl | 0.04 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0362 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.905 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-04-19 | ug/L | 0.865 | 2012-04-12 06:30 | 20121040014 | Accu | 210061400 | Accu 210061400: 20 µl | 0.04 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0346 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.865 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-06-08 | ug/L | 1.048 | 2013-05-30 08:00 | 20131510012 | Crescent | 04CC151 | Crescent 04CC151: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0524 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.048 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-05-22 | ug/L | 0.962 | 2013-05-16 07:30 | 20131370005 | Crescent, Accu | 04CC151, 213041120 | Crescent 04CC151: 10 µl & Accu 213041120: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0481 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.962 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2014-01-15 | ug/L | 0.844 | 2014-01-09 07:00 | 20140100001 | Crescent, Accu | 04CC151, 213041120 | Crescent 04CC151: 10 µl & Accu 213041120: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0422 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.844 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-11-15 | ug/L | 1.008 | 2013-11-14 08:00 | 20133190010 | Abs, Crescent | 051413, 04CC151 | Abs 051413: 10 µl & Crescent 04CC151: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0504 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.008 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-06-29 | ug/L | 1.108 | 2013-06-20 07:30 | 20131720017 | ChemService, Crescent | 1530700, 04CC151 | ChemService 1530700: 10 µl & Crescent 04CC151: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0554 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.108 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-09-27 | ug/L | 1.008 | 2013-09-25 06:00 | 20132690017 | Abs, NSI | 051413, 053013 | Abs 051413: 10 µl & NSI 053013: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0504 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.008 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2014-02-22 | ug/L | 1.062 | 2014-02-13 07:30 | 20140440053 | ChemService, NSI | 1530700, 053013 | ChemService 1530700: 10 µl & NSI 053013: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0531 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.062 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-09-17 | ug/L | 0.874 | 2013-09-06 06:30 | 20132490347 | Crescent, Supelco | 04CC151, LC00154 | Crescent 04CC151: 10 µl & Supelco LC00154: 10 µl | 0.05 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0437 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.874 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-10-14 | ug/L | 1 | 2010-10-07 07:00 | 20102810025 | Abs | 080510 | Abs 080510: 20 µl | 0.052 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.052 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-11-25 | ug/L | 0.77692 | 2010-11-18 07:00 | 20103230116 | Abs | 080510 | Abs 080510: 20 µl | 0.052 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0404 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77692 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-04-05 | ug/L | 1.0981 | 2011-03-31 07:00 | 20110910086 | Abs | 080510 | Abs 080510: 20 µl | 0.052 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0571 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0981 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-06-27 | ug/L | 0.99231 | 2011-06-23 06:30 | 20111750001 | Abs | 080510 | Abs 080510: 20 µl | 0.052 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0516 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99231 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-10-18 | ug/L | 1.2231 | 2011-10-07 06:45 | 20112840032 | Abs | 080510 | Abs 080510: 20 µl | 0.052 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0636 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2231 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-02-18 | ug/L | 1.0635 | 2012-02-02 12:00 | 20120370005 | Abs | 080510 | Abs 080510: 20 µl | 0.052 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0553 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0635 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-11-04 | ug/L | 1.31 | 2009-11-02 08:00 | 20093060179 | Spex Organics | C1090901006 | Spex Organics C1090901006: 10 µl | 0.06 | ETHYL_TERT-BUTYL_ETHER | E | 0.032 | 0.0786 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.31 | Y | Estimated Values | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-02-09 | ug/L | 1.0117 | 2010-02-04 06:00 | 20100360004 | Spex Organics | C1090901006 | Spex Organics C1090901006: 10 µl | 0.06 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0607 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0117 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-04-27 | ug/L | 0.97167 | 2010-04-22 06:00 | 20101130058 | Spex Organics | C1090901006 | Spex Organics C1090901006: 10 µl | 0.06 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0583 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97167 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-07-09 | ug/L | 1.09 | 2010-07-08 06:00 | 20101900011 | Spex Organics | C1090901006 | Spex Organics C1090901006: 10 µl | 0.06 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0654 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.09 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-11-02 | ug/L | 0.84375 | 2010-10-28 07:00 | 20103020157 | Accu | 210061400 | Accu 210061400: 40 µl | 0.08 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0675 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84375 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-03-17 | ug/L | 0.84625 | 2011-03-10 06:00 | 20110690081 | Accu | 210061400 | Accu 210061400: 40 µl | 0.08 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0677 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84625 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-06-05 | ug/L | 0.83 | 2011-06-02 07:00 | 20111540001 | Accu | 210061400 | Accu 210061400: 40 µl | 0.08 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0664 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-08-27 | ug/L | 0.96625 | 2011-08-25 09:00 | 20112380060 | Accu | 210061400 | Accu 210061400: 40 µl | 0.08 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0773 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96625 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-12-17 | ug/L | 0.96 | 2011-12-15 07:00 | 20113500023 | Accu | 210061400 | Accu 210061400: 40 µl | 0.08 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0768 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-12-12 | ug/L | 1.03 | 2013-11-26 07:30 | 20133310034 | NSI, Accu | 053013, 213041120 | NSI 053013: 20 µl & Accu 213041120: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.103 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.03 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-12-31 | ug/L | 1.04 | 2013-12-19 07:00 | 20133540012 | Abs, Crescent | 051413, 04CC151 | Abs 051413: 20 µl & Crescent 04CC151: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.104 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.04 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-08-22 | ug/L | 0.94 | 2013-08-15 07:00 | 20132280015 | ChemService, NSI | 1530700, 053013 | ChemService 1530700: 20 µl & NSI 053013: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.094 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2014-03-19 | ug/L | 1.09 | 2014-03-19 07:30 | 20140780042 | ChemService, NSI | 1530700, 053013 | ChemService 1530700: 20 µl & NSI 053013: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.109 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.09 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-09-17 | ug/L | 0.889 | 2013-08-29 07:00 | 20132420006 | Crescent, Supelco | 04CC151, LC00154 | Crescent 04CC151: 20 µl & Supelco LC00154: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0889 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.889 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-07-24 | ug/L | 1.04 | 2013-07-18 07:00 | 20132000017 | NSI, Supelco | 053013, LC00154 | NSI 053013: 20 µl & Supelco LC00154: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.104 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.04 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2014-03-01 | ug/L | 0.739 | 2014-02-27 07:30 | 20140590009 | NSI, Supelco | 053013, LC00154 | NSI 053013: 20 µl & Supelco LC00154: 20 µl | 0.1 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0739 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.739 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-02-02 | ug/L | 1.0192 | 2011-01-27 06:00 | 20110280003 | Abs | 080510 | Abs 080510: 40 µl | 0.104 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.106 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0192 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-05-17 | ug/L | 0.81346 | 2011-05-12 06:30 | 20111330146 | Abs | 080510 | Abs 080510: 40 µl | 0.104 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0846 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81346 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-08-13 | ug/L | 0.99038 | 2011-08-10 07:30 | 20112220124 | Abs | 080510 | Abs 080510: 40 µl | 0.104 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.103 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99038 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2011-11-22 | ug/L | 0.99038 | 2011-11-17 06:30 | 20113220201 | Abs | 080510 | Abs 080510: 40 µl | 0.104 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.103 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99038 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-03-24 | ug/L | 0.91923 | 2012-03-22 06:00 | 20120830082 | Abs | 080510 | Abs 080510: 40 µl | 0.104 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.0956 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91923 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-01-27 | ug/L | 0.91667 | 2009-01-15 08:00 | 20090160001 | ChemService | 406-110C | ChemService 406-110C: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-04-11 | ug/L | 0.91667 | 2009-04-09 09:00 | 20090990153 | ChemService | 406-110C | ChemService 406-110C: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-05-22 | ug/L | 0.83333 | 2009-05-20 08:00 | 20091410007 | ChemService | 406-110C | ChemService 406-110C: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | E | 0.04 | 0.1 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83333 | Y | Estimated Values | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-06-23 | ug/L | 0.91667 | 2009-06-18 08:30 | 20091700135 | ChemService | 406-110C | ChemService 406-110C: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-09-11 | ug/L | 0.75 | 2009-09-10 07:30 | 20092530076 | ChemService | 406-110C | ChemService 406-110C: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | E | 0.04 | 0.09 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75 | Y | Estimated Values | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-02-12 | ug/L | 1 | 2009-02-05 09:00 | 20090370001 | NSI | 081108-01VOC | NSI 081108-01VOC: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.12 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-06-06 | ug/L | 0.91667 | 2009-05-28 05:40 | 20091490010 | NSI | 081108-01VOC | NSI 081108-01VOC: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-07-09 | ug/L | 0.91667 | 2009-07-07 08:00 | 20091880141 | NSI | 081108-01VOC | NSI 081108-01VOC: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-08-27 | ug/L | 0.91667 | 2009-08-20 07:30 | 20092330007 | NSI | 081108-01VOC | NSI 081108-01VOC: 20 µl | 0.12 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-07-19 | ug/L | 0.92857 | 2012-07-12 07:30 | 20121950053 | Crescent, Accu | 05BB048, 212051193 | Crescent 05BB048: 10 µl & Accu 212051193: 10 µl | 0.14 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.13 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.92857 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-10-11 | ug/L | 0.93571 | 2012-10-04 06:30 | 20122790029 | Crescent, ChemService | 05BB048, 472200 | Crescent 05BB048: 10 µl & ChemService 472200: 10 µl | 0.14 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.131 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93571 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-02-28 | ug/L | 0.91333 | 2013-02-28 07:00 | 20130590063 | NSI, Accu | 062112-01VOA, 212051193 | NSI 062112-01VOA: 10 µl & Accu 212051193: 10 µl | 0.15 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.137 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91333 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-12-09 | ug/L | 0.87333 | 2012-12-06 06:00 | 20123420073 | NSI, ChemService | 062112-01VOA, 472200 | NSI 062112-01VOA: 10 µl & ChemService 472200: 10 µl | 0.15 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.131 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.87333 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-12-15 | ug/L | 0.93889 | 2009-12-10 08:00 | 20093450027 | Spex Organics | C1090901006 | Spex Organics C1090901006: 30 µl | 0.18 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.169 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93889 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-03-17 | ug/L | 0.84444 | 2010-03-11 06:00 | 20100710179 | Spex Organics | C1090901006 | Spex Organics C1090901006: 30 µl | 0.18 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.152 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84444 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-06-09 | ug/L | 0.99444 | 2010-06-03 06:00 | 20101550002 | Spex Organics | C1090901006 | Spex Organics C1090901006: 30 µl | 0.18 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.179 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99444 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2010-08-21 | ug/L | 0.93333 | 2010-08-19 06:00 | 20102320014 | Spex Organics | C1090901006 | Spex Organics C1090901006: 30 µl | 0.18 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.168 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93333 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-02-27 | ug/L | 0.91667 | 2009-02-26 08:00 | 20090580002 | ChemService | 406-110C | ChemService 406-110C: 40 µl | 0.24 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.22 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91667 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-08-04 | ug/L | 0.875 | 2009-07-30 08:00 | 20092110201 | ChemService | 406-110C | ChemService 406-110C: 40 µl | 0.24 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.21 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.875 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-04-08 | ug/L | 0.875 | 2009-03-19 08:00 | 20090780140 | NSI | 081108-01VOC | NSI 081108-01VOC: 40 µl | 0.24 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.21 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.875 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2009-05-06 | ug/L | 1 | 2009-04-30 07:00 | 20091210008 | NSI | 081108-01VOC | NSI 081108-01VOC: 40 µl | 0.24 | ETHYL_TERT-BUTYL_ETHER | . | 0.04 | 0.24 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-11-08 | ug/L | 0.91786 | 2012-11-01 06:30 | 20123070013 | Crescent, Accu | 05BB048, 212051193 | Crescent 05BB048: 20 µl & Accu 212051193: 20 µl | 0.28 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.257 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91786 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-04-17 | ug/L | 0.80714 | 2013-04-11 06:30 | 20131020084 | Accu, Crescent | 212051389, 05BB048 | Accu 212051389: 20 µl & Crescent 05BB048: 20 µl | 0.28 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.226 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80714 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-08-28 | ug/L | 0.76333 | 2012-08-23 06:30 | 20122370105 | NSI, ChemService | 062112-01VOA, 472200 | NSI 062112-01VOA: 20 µl & ChemService 472200: 20 µl | 0.3 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.229 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.76333 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-09-11 | ug/L | 0.83333 | 2012-09-06 07:00 | 20122500213 | NSI, Supelco | 062112-01VOA, LB92787 | NSI 062112-01VOA: 20 µl & Supelco LB92787: 20 µl | 0.3 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.25 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83333 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-10-23 | ug/L | 1.0643 | 2012-10-18 06:30 | 20122930154 | Crescent, Abs | 05BB048, 051712 | Crescent 05BB048: 40 µl & Abs 051712: 40 µl | 0.56 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.596 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0643 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2013-02-14 | ug/L | 0.95 | 2013-02-07 06:00 | 20130390107 | Crescent, ChemService | 05BB048, 472200 | Crescent 05BB048: 40 µl & ChemService 472200: 40 µl | 0.56 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.532 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571
ETHYL TERT-BUTYL ETHER | 2021 | 50004GCM66 | 2012-12-22 | ug/L | 0.88833 | 2012-12-20 06:00 | 20123560093 | NSI, Supelco | 062112-01VOA, LB92787 | NSI 062112-01VOA: 40 µl & Supelco LB92787: 40 µl | 0.6 | ETHYL_TERT-BUTYL_ETHER | . | 0.032 | 0.533 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88833 | Y | Recoverable Result | Y |  | 69 | 0.94655 | 0.10631 | 0.93571 | 0.133 | 0.098592 | 0.93571