## OBSP Data
AnalyteName | Schedulecode | Parm+Method Code | Analysis Date | ReportUnit | RECOVERY | Login Date | Laboratory ID | ManufShortName | CALotNum | SpikeMix | Expected Conc. | newanalyte | Remark Code | REPORTING LEVEL | Lab Value | UNITS | All | Comments | Spike Level | Pct Recovery | Sample Spiked | STATUS | PLOT | perm_e | N | Mean | Std-Dev. | Median | IQR | F_Pseudo | Median Pct Recovery | maxout
TEBUTHIURON | 2033 | 82670GCM35 | 2009-02-11 | ug/L |  | 2009-01-08 07:00 | 20090080048 | Abs | 080408Pest | Abs 080408Pest: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 98 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-03-25 | ug/L |  | 2009-03-12 07:30 | 20090710012 | Abs | 080408Pest | Abs 080408Pest: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-06-30 | ug/L |  | 2009-05-20 07:30 | 20091410001 | Abs | 080408Pest | Abs 080408Pest: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-10-27 | ug/L |  | 2009-07-23 07:30 | 20092040202 | Abs | 080408Pest | Abs 080408Pest: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-01-10 | ug/L |  | 2009-08-27 07:30 | 20092400050 | Abs | 080408Pest | Abs 080408Pest: 25 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-07 | ug/L |  | 2010-10-07 07:30 | 20102810019 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-28 | ug/L |  | 2010-11-24 07:00 | 20103330007 | Abs | 080610 | Abs 080610: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-22 | ug/L |  | 2011-02-02 06:30 | 20110340007 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-03-16 | ug/L |  | 2011-03-03 07:00 | 20110620082 | Abs | 080610 | Abs 080610: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-05-03 | ug/L |  | 2011-04-22 06:20 | 20111170107 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-05-29 | ug/L |  | 2011-05-19 07:30 | 20111390178 | Abs | 080610 | Abs 080610: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-08-22 | ug/L |  | 2011-07-28 07:00 | 20112100007 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-10-20 | ug/L |  | 2011-10-07 06:30 | 20112840031 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-01-20 | ug/L |  | 2012-01-12 07:00 | 20120120071 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-02-09 | ug/L |  | 2012-01-26 07:00 | 20120260028 | Abs | 080610 | Abs 080610: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-03-08 | ug/L |  | 2012-02-23 07:30 | 20120540036 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-04-12 | ug/L |  | 2012-03-29 08:00 | 20120930003 | Abs | 080610 | Abs 080610: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-04-24 | ug/L |  | 2012-04-12 07:00 | 20121040012 | Abs | 080610 | Abs 080610: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-02-23 | ug/L |  | 2009-11-19 08:00 | 20093240038 | Accu | 209081187 | Accu 209081187: 100 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-03-24 | ug/L |  | 2010-01-21 07:00 | 20100210100 | Accu | 209081187 | Accu 209081187: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-05-07 | ug/L |  | 2010-02-25 06:30 | 20100570086 | Accu | 209081187 | Accu 209081187: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-09-09 | ug/L |  | 2010-06-24 06:30 | 20101760006 | Accu | 209081187 | Accu 209081187: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-12-14 | ug/L |  | 2010-09-02 06:30 | 20102500051 | Accu | 209081187 | Accu 209081187: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-12 | ug/L |  | 2010-10-28 07:30 | 20103020155 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-01 | ug/L |  | 2010-12-16 06:30 | 20103510060 | Accu | 210061341 | Accu 210061341: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-14 | ug/L |  | 2011-01-20 07:30 | 20110210095 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-04-07 | ug/L |  | 2011-03-24 06:30 | 20110840108 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-05-21 | ug/L |  | 2011-04-28 07:00 | 20111190184 | Accu | 210061341 | Accu 210061341: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-07-02 | ug/L |  | 2011-06-16 08:00 | 20111680035 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-07-20 | ug/L |  | 2011-07-07 07:00 | 20111890004 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-08-30 | ug/L |  | 2011-08-04 08:00 | 20112170174 | Accu | 210061341 | Accu 210061341: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-11-05 | ug/L |  | 2011-10-27 07:30 | 20113010047 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-12-08 | ug/L |  | 2011-11-29 07:00 | 20113340046 | Accu | 210061341 | Accu 210061341: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-03-02 | ug/L |  | 2012-02-09 07:00 | 20120400058 | Accu | 210061341 | Accu 210061341: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-03-27 | ug/L |  | 2012-03-15 08:00 | 20120760111 | Accu | 210061341 | Accu 210061341: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-03-03 | ug/L |  | 2009-02-05 08:30 | 20090370005 | Accu | B8090098 | Accu B8090098: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-05-13 | ug/L |  | 2009-04-16 07:30 | 20091060132 | Accu | B8090098 | Accu B8090098: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-09-03 | ug/L |  | 2009-06-25 07:30 | 20091770120 | Accu | B8090098 | Accu B8090098: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-12-30 | ug/L |  | 2009-08-20 07:00 | 20092330002 | Accu | B8090098 | Accu B8090098: 25 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-02-18 | ug/L |  | 2009-01-21 07:00 | 20090210033 | ChemService | 406-99B | ChemService 406-99B: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-04-30 | ug/L |  | 2009-04-02 06:30 | 20090920067 | ChemService | 406-99B | ChemService 406-99B: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-08-12 | ug/L |  | 2009-06-11 08:00 | 20091620170 | ChemService | 406-99B | ChemService 406-99B: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-12-01 | ug/L |  | 2009-08-06 07:00 | 20092180180 | ChemService | 406-99B | ChemService 406-99B: 25 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-03-03 | ug/L |  | 2009-12-03 08:30 | 20093380009 | ChemService | 430-36C | ChemService 430-36C: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-06-26 | ug/L |  | 2010-04-22 06:30 | 20101130055 | ChemService | 430-36C | ChemService 430-36C: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-09-21 | ug/L |  | 2010-07-01 06:30 | 20101830110 | ChemService | 430-36C | ChemService 430-36C: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-06-03 | ug/L |  | 2009-04-30 06:30 | 20091210002 | Crescent | 08HH019 | Crescent 08HH019: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-07-23 | ug/L |  | 2009-06-04 07:30 | 20091550184 | Crescent | 08HH019 | Crescent 08HH019: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-06-03 | ug/L |  | 2010-04-01 06:30 | 20100920048 | Crescent | 08Ll064 | Crescent 08Ll064: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-07-21 | ug/L |  | 2010-05-20 06:30 | 20101410006 | Crescent | 08Ll064 | Crescent 08Ll064: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-10-19 | ug/L |  | 2010-08-03 06:30 | 20102160002 | Crescent | 08Ll064 | Crescent 08Ll064: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-09-13 | ug/L |  | 2011-08-18 08:47 | 20112310030 | ERA | 0318-13-02.2 | ERA 0318-13-02.2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-10-12 | ug/L |  | 2012-09-27 07:00 | 20122720048 | ERA | 0507-12-04.2 | ERA 0507-12-04.2: 50 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-03-12 | ug/L |  | 2009-02-26 07:30 | 20090580003 | ERA | 0804-08-02PEST | ERA 0804-08-02PEST: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-06-18 | ug/L |  | 2009-05-06 07:30 | 20091320001 | ERA | 0804-08-02PEST | ERA 0804-08-02PEST: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-09-16 | ug/L |  | 2009-07-07 07:30 | 20091880135 | ERA | 0804-08-02PEST | ERA 0804-08-02PEST: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-11-11 | ug/L |  | 2009-07-30 07:30 | 20092110204 | ERA | 0804-08-02PEST | ERA 0804-08-02PEST: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-01-17 | ug/L |  | 2009-09-03 06:30 | 20092460258 | ERA | 0804-08-02PEST | ERA 0804-08-02PEST: 25 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-01-21 | ug/L |  | 2009-09-10 07:00 | 20092530080 | ERA | 0804-08-02PEST | ERA 0804-08-02PEST: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-13 | ug/L |  | 2010-11-04 07:30 | 20103080116 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-08 | ug/L |  | 2011-01-06 07:00 | 20110070109 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-28 | ug/L |  | 2011-02-10 07:00 | 20110410061 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-04-13 | ug/L |  | 2011-03-31 07:30 | 20110910084 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-06-20 | ug/L |  | 2011-05-26 07:30 | 20111470065 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-07-28 | ug/L |  | 2011-07-14 08:00 | 20111960039 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-10-25 | ug/L |  | 2011-10-13 07:30 | 20112870073 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-12-03 | ug/L |  | 2011-11-17 07:00 | 20113220198 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-03-05 | ug/L |  | 2012-02-16 07:30 | 20120470078 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-05-08 | ug/L |  | 2012-04-26 07:00 | 20121180062 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-02-16 | ug/L |  | 2009-12-17 07:00 | 20093520041 | NSI | 090109-01Pest | NSI 090109-01Pest: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-03-25 | ug/L |  | 2010-01-28 06:30 | 20100280077 | NSI | 090109-01Pest | NSI 090109-01Pest: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-08-03 | ug/L |  | 2010-06-03 06:30 | 20101550006 | NSI | 090109-01Pest | NSI 090109-01Pest: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-10-27 | ug/L |  | 2010-08-12 06:30 | 20102250096 | NSI | 090109-01Pest | NSI 090109-01Pest: 100 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-02-11 | ug/L |  | 2009-01-15 07:30 | 20090160003 | Protocol | D1080908018 | Protocol D1080908018: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-04-28 | ug/L |  | 2009-03-26 07:30 | 20090860072 | Protocol | D1080908018 | Protocol D1080908018: 25 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-07-10 | ug/L |  | 2009-05-14 09:00 | 20091340062 | Protocol | D1080908018 | Protocol D1080908018: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-09-15 | ug/L |  | 2009-07-02 08:00 | 20091830169 | Protocol | D1080908018 | Protocol D1080908018: 40 µl |  | TEBUTHIURON | < | 0.02 | 0.02 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-02-02 | ug/L |  | 2009-10-08 08:30 | 20092820002 | Restek | A069639 | Restek A069639: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-03-19 | ug/L |  | 2010-01-14 06:30 | 20100150090 | Restek | A069639 | Restek A069639: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-07-29 | ug/L |  | 2010-05-27 06:30 | 20101480005 | Restek | A069639 | Restek A069639: 100 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-10-18 | ug/L |  | 2010-08-05 06:30 | 20102180092 | Restek | A069639 | Restek A069639: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-06-28 | ug/L |  | 2012-06-21 06:00 | 20121730123 | Restek | A088380 | Restek A088380: 50 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-11 | ug/L |  | 2010-10-14 07:30 | 20102880001 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-21 | ug/L |  | 2010-11-30 07:00 | 20103350003 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-19 | ug/L |  | 2011-01-27 06:30 | 20110280004 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-03-23 | ug/L |  | 2011-03-10 06:30 | 20110690078 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-05-14 | ug/L |  | 2011-05-05 07:30 | 20111260154 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-07-07 | ug/L |  | 2011-06-23 07:00 | 20111750006 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-09-19 | ug/L |  | 2011-08-25 10:00 | 20112380055 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-09-29 | ug/L |  | 2011-09-08 07:30 | 20112520209 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-11-17 | ug/L |  | 2011-11-03 07:00 | 20113080001 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-12-21 | ug/L |  | 2011-12-13 07:30 | 20113490056 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-01-27 | ug/L |  | 2012-01-19 07:00 | 20120200073 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-04-05 | ug/L |  | 2012-03-22 06:30 | 20120830077 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-06-02 | ug/L |  | 2012-05-24 07:00 | 20121460001 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-02-06 | ug/L |  | 2009-10-22 08:30 | 20092960003 | Supelco | LB69639 | Supelco LB69639: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-04-22 | ug/L |  | 2010-02-11 06:30 | 20100420055 | Supelco | LB69639 | Supelco LB69639: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-07-07 | ug/L |  | 2010-05-06 06:30 | 20101270038 | Supelco | LB69639 | Supelco LB69639: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-09-23 | ug/L |  | 2010-07-13 06:30 | 20101950137 | Supelco | LB69639 | Supelco LB69639: 40 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-04-29 | ug/L |  | 2014-04-17 06:00 | 20141080004 | Crescent, Abs | 03DD006, 032714 | Crescent 03DD006: 30 µl & Abs 032714: 30 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-08-27 | ug/L |  | 2014-08-14 06:00 | 20142270001 | Crescent, Abs | 03DD006, 032714 | Crescent 03DD006: 60 µl & Abs 032714: 60 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-05-20 | ug/L |  | 2014-05-01 06:00 | 20141220025 | Supelco, Abs | LC06435, 032714 | Supelco LC06435: 60 µl & Abs 032714: 60 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-11-19 | ug/L |  | 2014-11-06 06:00 | 20143110108 | Spex Organics, Crescent | E1140407002, 03DD006 | Spex Organics E1140407002: 60 µl & Crescent 03DD006: 60 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-07-11 | ug/L |  | 2013-06-27 06:30 | 20131790009 | ChemService, ERA | 1473000, 0318-13-02.2 | ChemService 1473000: 20 µl & ERA 0318-13-02.2: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-09-20 | ug/L |  | 2013-09-06 06:50 | 20132490342 | Spex Organics, ERA | T1130514013, 0318-13-02.2 | Spex Organics T1130514013: 100 µl & ERA 0318-13-02.2: 100 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-02-12 | ug/L |  | 2014-02-06 06:30 | 20140370077 | Spex Organics, ERA | T1130514013, 0318-13-02.2 | Spex Organics T1130514013: 50 µl & ERA 0318-13-02.2: 50 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-09-05 | ug/L |  | 2012-08-16 06:30 | 20122300021 | Spex Organics, ERA | T1120530013, 0507-12-04.2 | Spex Organics T1120530013: 20 µl & ERA 0507-12-04.2: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-01-15 | ug/L |  | 2012-12-20 06:30 | 20123560092 | Spex Organics, ERA | T1120530013, 0507-12-04.2 | Spex Organics T1120530013: 50 µl & ERA 0507-12-04.2: 50 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-06-30 | ug/L |  | 2014-06-12 06:00 | 20141640076 | Crescent, Spex Organics | 03DD006, E1140407002 | Crescent 03DD006: 30 µl & Spex Organics E1140407002: 30 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-07-17 | ug/L |  | 2014-07-10 06:00 | 20141910243 | Supelco, Spex Organics | LC06435, E1140407002 | Supelco LC06435: 30 µl & Spex Organics E1140407002: 30 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-10-04 | ug/L |  | 2014-09-25 06:00 | 20142720002 | Supelco, Spex Organics | LC06435, E1140407002 | Supelco LC06435: 60 µl & Spex Organics E1140407002: 60 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-07-14 | ug/L |  | 2012-07-03 08:00 | 20121870008 | Restek, Supelco | A088380, LB92783 | Restek A088380: 20 µl & Supelco LB92783: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-11-20 | ug/L |  | 2012-11-08 06:30 | 20123140095 | Spex Organics, Supelco | T1120530013, LB92783 | Spex Organics T1120530013: 100 µl & Supelco LB92783: 100 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-02-22 | ug/L |  | 2013-02-07 06:30 | 20130390105 | Spex Organics, Supelco | T1120530013, LB92783 | Spex Organics T1120530013: 20 µl & Supelco LB92783: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-04-02 | ug/L |  | 2013-03-21 06:30 | 20130810108 | Spex Organics, Supelco | T1120530013, LB92783 | Spex Organics T1120530013: 50 µl & Supelco LB92783: 50 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-06-05 | ug/L |  | 2013-05-16 07:00 | 20131370001 | ChemService, Supelco | 1473000, LB99392 | ChemService 1473000: 50 µl & Supelco LB99392: 50 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-01-08 | ug/L |  | 2013-12-12 06:30 | 20133470008 | ChemService, Supelco | 1473000, LB99392 | ChemService 1473000: 20 µl & Supelco LB99392: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-11-09 | ug/L |  | 2013-10-31 06:30 | 20133050050 | Spex Organics, Supelco | T1130514013, LB99392 | Spex Organics T1130514013: 20 µl & Supelco LB99392: 20 µl |  | TEBUTHIURON | < | 0.028 | 0.028 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N |  | 0 | [5 empty] | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-03-05 | ug/L | 2.1091 | 2011-02-17 08:30 | 20110490035 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 25 µl | 0.02475 | TEBUTHIURON | E | 0.028 | 0.0522 | MICROGRAMS PER LITER |  |  | < Reporting Level | 2 | Y | Estimated Values | Y |  | 1 | 2 |  | 2 | 0 | 0 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-06-30 | ug/L | 2.1969 | 2010-04-29 10:30 | 20101190091 | NWQL | PSP2SP116C | NWQL PSP2SP116C: 40 µl | 0.038782 | TEBUTHIURON | . | 0.028 | 0.0852 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 2 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-12-16 | ug/L | 1.8428 | 2010-09-09 06:40 | 20102530008 | NWQL | PSP2SP116C | NWQL PSP2SP116C: 40 µl | 0.040428 | TEBUTHIURON | . | 0.028 | 0.0745 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.8428 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-08-25 | ug/L | 1.6844 | 2010-06-17 10:30 | 20101720149 | NWQL | PSP2SP116C | NWQL PSP2SP116C: 40 µl | 0.040547 | TEBUTHIURON | . | 0.028 | 0.0683 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6844 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-08-16 | ug/L | 1.4176 | 2014-08-07 07:00 | 20142200164 | NSI, Crescent | 042814, 03DD006 | NSI 042814: 30 µl & Crescent 03DD006: 30 µl | 0.0558 | TEBUTHIURON | . | 0.028 | 0.0791 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.4176 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-10-21 | ug/L | 1.2079 | 2014-10-09 06:00 | 20142820151 | Abs, NSI | 032714, 042814 | Abs 032714: 30 µl & NSI 042814: 30 µl | 0.0558 | TEBUTHIURON | . | 0.028 | 0.0674 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2079 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-09-22 | ug/L | 1.2133 | 2014-09-03 06:00 | 20142470041 | Spex Organics, NSI | E1140407002, 042814 | Spex Organics E1140407002: 30 µl & NSI 042814: 30 µl | 0.0558 | TEBUTHIURON | . | 0.028 | 0.0677 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2133 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-12-15 | ug/L | 1.1434 | 2014-12-03 06:00 | 20143370094 | NSI, Supelco | 042814, LC06435 | NSI 042814: 30 µl & Supelco LC06435: 30 µl | 0.0558 | TEBUTHIURON | . | 0.028 | 0.0638 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1434 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-06-24 | ug/L | 0.64528 | 2014-05-22 07:00 | 20141420140 | ERA, Crescent | 0227-14-08, 03DD006 | ERA 0227-14-08: 30 µl & Crescent 03DD006: 30 µl | 0.0561 | TEBUTHIURON | . | 0.028 | 0.0362 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.64528 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-06-26 | ug/L | 1.5169 | 2014-05-29 06:00 | 20141490205 | Abs, ERA | 032714, 0227-14-08 | Abs 032714: 30 µl & ERA 0227-14-08: 30 µl | 0.0561 | TEBUTHIURON | . | 0.028 | 0.0851 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5169 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-05-28 | ug/L | 0.96257 | 2014-05-16 06:00 | 20141360094 | Spex Organics, ERA | E1140407002, 0227-14-08 | Spex Organics E1140407002: 30 µl & ERA 0227-14-08: 30 µl | 0.0561 | TEBUTHIURON | . | 0.028 | 0.054 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96257 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-07-16 | ug/L | 1.0642 | 2014-07-02 07:00 | 20141840097 | ERA, Supelco | 0227-14-08, LC06435 | ERA 0227-14-08: 30 µl & Supelco LC06435: 30 µl | 0.0561 | TEBUTHIURON | . | 0.028 | 0.0597 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0642 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-05-22 | ug/L | 1.7391 | 2013-05-02 06:00 | 20131230084 | Restek, Crescent | A088380, 05BB029 | Restek A088380: 20 µl & Crescent 05BB029: 20 µl | 0.0598 | TEBUTHIURON | . | 0.028 | 0.104 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.7391 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-08-29 | ug/L | 1.1957 | 2012-08-07 06:00 | 20122210006 | Spex Organics, Crescent | T1120530013, 05BB029 | Spex Organics T1120530013: 20 µl & Crescent 05BB029: 20 µl | 0.0598 | TEBUTHIURON | E | 0.028 | 0.0715 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1957 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-09-17 | ug/L | 1.5017 | 2012-08-30 06:30 | 20122430227 | Crescent, ERA | 05BB029, 0507-12-04.2 | Crescent 05BB029: 20 µl & ERA 0507-12-04.2: 20 µl | 0.0598 | TEBUTHIURON | . | 0.028 | 0.0898 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5017 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-03-07 | ug/L | 1.0619 | 2013-02-19 06:30 | 20130510051 | Crescent, ERA | 05BB029, 0507-12-04.2 | Crescent 05BB029: 20 µl & ERA 0507-12-04.2: 20 µl | 0.0598 | TEBUTHIURON | . | 0.028 | 0.0635 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0619 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-06-16 | ug/L | 1.5314 | 2012-06-07 06:30 | 20121600156 | Abs | 052412 | Abs 052412: 20 µl | 0.0606 | TEBUTHIURON | . | 0.028 | 0.0928 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5314 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-07-23 | ug/L | 2.0462 | 2013-07-11 06:30 | 20131920149 | Restek, Abs | A093651, 012513 | Restek A093651: 20 µl & Abs 012513: 20 µl | 0.0606 | TEBUTHIURON | E | 0.028 | 0.124 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 2 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-09-26 | ug/L | 1.0132 | 2012-09-13 06:00 | 20122570122 | Restek, Abs | A088380, 052412 | Restek A088380: 20 µl & Abs 052412: 20 µl | 0.0606 | TEBUTHIURON | . | 0.028 | 0.0614 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0132 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-03-18 | ug/L | 1.1386 | 2013-03-05 06:00 | 20130650005 | Restek, Abs | A088380, 052412 | Restek A088380: 20 µl & Abs 052412: 20 µl | 0.0606 | TEBUTHIURON | . | 0.028 | 0.069 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1386 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-08-14 | ug/L | 2.2937 | 2012-07-19 06:30 | 20122020004 | Abs, Supelco | 052412, LB92783 | Abs 052412: 20 µl & Supelco LB92783: 20 µl | 0.0606 | TEBUTHIURON | E | 0.028 | 0.139 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 2 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-04-15 | ug/L | 1.2706 | 2014-03-27 06:30 | 20140860102 | Restek, Abs, Crescent | A093651, 012513, 10802AL | Restek A093651: 20 µl & Abs 012513: 20 µl & Crescent 10802AL: 10 µl | 0.0606 | TEBUTHIURON | . | 0.028 | 0.077 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2706 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-11-04 | ug/L | 1.505 | 2013-10-17 06:30 | 20132910008 | Restek, Abs, Crescent | A093651, 052412, 10802AL | Restek A093651: 20 µl & Abs 052412: 20 µl & Crescent 10802AL: 5 µl | 0.0606 | TEBUTHIURON | . | 0.028 | 0.0912 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.505 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-01-08 | ug/L | 1.071 | 2013-12-05 08:00 | 20133400002 | Restek, Abs, Crescent | A093651, 052412, 10802AL | Restek A093651: 20 µl & Abs 052412: 20 µl & Crescent 10802AL: 5 µl | 0.0606 | TEBUTHIURON | . | 0.028 | 0.0649 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.071 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-09 | ug/L | 1.6834 | 2010-10-21 07:30 | 20102950072 | Restek | A075606 | Restek A075606: 40 µl | 0.065344 | TEBUTHIURON | . | 0.028 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6834 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-18 | ug/L | 1.537 | 2010-11-10 07:30 | 20103140135 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.067012 | TEBUTHIURON | . | 0.028 | 0.103 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.537 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-03-30 | ug/L | 1.2865 | 2011-03-17 08:00 | 20110770001 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.06716 | TEBUTHIURON | . | 0.028 | 0.0864 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2865 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-08-17 | ug/L | 1.4122 | 2011-07-21 07:30 | 20112030057 | Restek | A075606 | Restek A075606: 40 µl | 0.0672 | TEBUTHIURON | E | 0.028 | 0.0949 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.4122 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-12-04 | ug/L | 1.5327 | 2011-11-22 07:30 | 20113260111 | Restek | A075606 | Restek A075606: 40 µl | 0.0672 | TEBUTHIURON | . | 0.028 | 0.103 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5327 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-02-14 | ug/L | 1.6369 | 2012-02-02 12:10 | 20120370001 | Restek | A075606 | Restek A075606: 40 µl | 0.0672 | TEBUTHIURON | E | 0.028 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6369 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-07-20 | ug/L | 1.6667 | 2011-07-07 06:30 | 20111890005 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.0672 | TEBUTHIURON | E | 0.028 | 0.112 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6667 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-09-24 | ug/L | 1.2783 | 2011-09-01 07:30 | 20112450022 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.0672 | TEBUTHIURON | . | 0.028 | 0.0859 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2783 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-12-22 | ug/L | 1.6815 | 2011-12-15 07:30 | 20113500021 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.0672 | TEBUTHIURON | E | 0.028 | 0.113 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6815 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-03-09 | ug/L | 1.875 | 2012-03-01 07:30 | 20120620015 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.0672 | TEBUTHIURON | . | 0.028 | 0.126 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.875 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-04-23 | ug/L | 1.7094 | 2011-04-14 07:30 | 20111050150 | Restek | A075606 | Restek A075606: 40 µl | 0.067858 | TEBUTHIURON | E | 0.028 | 0.116 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.7094 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-05-25 | ug/L | 1.77 | 2011-05-12 07:00 | 20111330148 | Supelco | LB77168 | Supelco LB77168: 40 µl | 0.068362 | TEBUTHIURON | . | 0.028 | 0.121 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.77 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-03-08 | ug/L | 1.5347 | 2011-02-24 08:00 | 20110560005 | Restek | A075606 | Restek A075606: 40 µl | 0.068418 | TEBUTHIURON | E | 0.028 | 0.105 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5347 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-06-22 | ug/L | 1.849 | 2011-06-07 10:00 | 20111590005 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 75 µl | 0.074096 | TEBUTHIURON | E | 0.028 | 0.137 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.849 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-03-04 | ug/L | 1.2538 | 2009-02-12 06:30 | 20090430172 | NWQL | PSP2SP115A | NWQL PSP2SP115A: 75 µl | 0.074177 | TEBUTHIURON | . | 0.02 | 0.093 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2538 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-09-03 | ug/L | 1.384 | 2009-06-25 08:00 | 20091770121 | NWQL | PSP2SP115C | NWQL PSP2SP115C: 75 µl | 0.074419 | TEBUTHIURON | . | 0.02 | 0.103 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.384 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-07-13 | ug/L | 1.0289 | 2009-05-28 05:30 | 20091490007 | NWQL | PSP2SP115C | NWQL PSP2SP115C: 75 µl | 0.074835 | TEBUTHIURON | . | 0.02 | 0.077 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0289 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-04-17 | ug/L | 1.663 | 2009-03-19 07:30 | 20090780142 | NWQL | PSP2SP115C | NWQL PSP2SP115C: 75 µl | 0.075165 | TEBUTHIURON | . | 0.02 | 0.125 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.663 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-05-27 | ug/L | 1.3452 | 2009-04-23 07:30 | 20091130130 | NWQL | PSP2SP115C | NWQL PSP2SP115C: 75 µl | 0.075827 | TEBUTHIURON | . | 0.02 | 0.102 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3452 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-06-11 | ug/L | 1.3191 | 2013-05-30 07:30 | 20131510026 | ERA, Crescent | 0318-13-02.2, 04CC101 | ERA 0318-13-02.2: 20 µl & Crescent 04CC101: 20 µl | 0.0796 | TEBUTHIURON | . | 0.028 | 0.105 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3191 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-09-05 | ug/L | 1.6332 | 2013-08-15 06:30 | 20132280016 | Crescent, Spex Organics | 04CC101, T1130514013 | Crescent 04CC101: 20 µl & Spex Organics T1130514013: 20 µl | 0.0796 | TEBUTHIURON | . | 0.028 | 0.13 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6332 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-12-05 | ug/L | 1.2023 | 2013-11-26 07:00 | 20133310030 | Crescent, ChemService, Cresc | 04CC101, 1473000, 10802AL | Crescent 04CC101: 20 µl & ChemService 1473000: 20 µl & Crescent 10802AL: 20 µl | 0.0796 | TEBUTHIURON | . | 0.028 | 0.0957 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2023 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-04-03 | ug/L | 0.71985 | 2014-03-13 07:00 | 20140730138 | Crescent, ChemService, Cresc | 04CC101, 1473000, 10802AL | Crescent 04CC101: 20 µl & ChemService 1473000: 20 µl & Crescent 10802AL: 10 µl | 0.0796 | TEBUTHIURON | . | 0.028 | 0.0573 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.71985 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-03-05 | ug/L | 0.80402 | 2014-02-20 06:30 | 20140510043 | Supelco, Crescent, Crescent | LB99392, 04CC101, 10802AL | Supelco LB99392: 20 µl & Crescent 04CC101: 20 µl & Crescent 10802AL: 10 µl | 0.0796 | TEBUTHIURON | . | 0.028 | 0.064 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80402 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-06-26 | ug/L | 1.9576 | 2013-06-13 06:30 | 20131650073 | Abs, ChemService | 042413, 1473000 | Abs 042413: 20 µl & ChemService 1473000: 20 µl | 0.0802 | TEBUTHIURON | . | 0.028 | 0.157 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.9576 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-01-10 | ug/L | 0.72195 | 2013-12-23 06:30 | 20133580036 | ERA, Abs, Crescent | 0318-13-02.2, 042413, 10802AL | ERA 0318-13-02.2: 20 µl & Abs 042413: 20 µl & Crescent 10802AL: 10 µl | 0.0802 | TEBUTHIURON | . | 0.028 | 0.0579 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72195 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-11-23 | ug/L | 1.202 | 2013-11-14 07:30 | 20133190012 | Supelco, Abs, Crescent | LB99392, 042413, 10802AL | Supelco LB99392: 20 µl & Abs 042413: 20 µl & Crescent 10802AL: 10 µl | 0.0802 | TEBUTHIURON | . | 0.028 | 0.0964 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.202 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-03-10 | ug/L | 1.2874 | 2009-02-19 06:45 | 20090500153 | NWQL | PSP2SP115A | NWQL PSP2SP115A: 75 µl | 0.083112 | TEBUTHIURON | . | 0.02 | 0.107 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2874 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-04-20 | ug/L | 1.6887 | 2011-04-07 09:00 | 20110980030 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 100 µl | 0.099483 | TEBUTHIURON | . | 0.028 | 0.168 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6887 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-20 | ug/L | 1.6782 | 2010-11-18 09:00 | 20103230114 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 100 µl | 0.1007 | TEBUTHIURON | E | 0.028 | 0.169 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6782 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-06-16 | ug/L | 1.5574 | 2010-04-08 06:30 | 20100990014 | Spex Organics | T1090901009 | Spex Organics T1090901009: 20 µl | 0.11044 | TEBUTHIURON | . | 0.028 | 0.172 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5574 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-12-01 | ug/L | 1.0889 | 2010-08-26 06:30 | 20102390088 | Ultra | CF-3558 | Ultra CF-3558: 20 µl | 0.11112 | TEBUTHIURON | E | 0.028 | 0.121 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0889 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-10-30 | ug/L | 0.94982 | 2014-10-23 06:00 | 20142960086 | NSI, Crescent | 042814, 03DD006 | NSI 042814: 60 µl & Crescent 03DD006: 60 µl | 0.1116 | TEBUTHIURON | . | 0.028 | 0.106 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94982 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-12-04 | ug/L | 1.2545 | 2014-11-19 06:00 | 20143240095 | Abs, NSI | 032714, 042814 | Abs 032714: 60 µl & NSI 042814: 60 µl | 0.1116 | TEBUTHIURON | . | 0.028 | 0.14 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2545 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-09-23 | ug/L | 0.78136 | 2014-09-10 06:00 | 20142530087 | NSI, Supelco | 042814, LC06435 | NSI 042814: 60 µl & Supelco LC06435: 60 µl | 0.1116 | TEBUTHIURON | . | 0.028 | 0.0872 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78136 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-05-19 | ug/L | 2.8908 | 2010-03-11 06:30 | 20100710178 | Ultra | CF-3558 | Ultra CF-3558: 20 µl | 0.11173 | TEBUTHIURON | . | 0.028 | 0.323 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 2 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-07-09 | ug/L | 1.0071 | 2014-06-26 06:00 | 20141770149 | Spex Organics, ERA | E1140407002, 0227-14-08 | Spex Organics E1140407002: 60 µl & ERA 0227-14-08: 60 µl | 0.1122 | TEBUTHIURON | E | 0.028 | 0.113 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0071 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-06-28 | ug/L | 1.7291 | 2014-06-05 07:00 | 20141570088 | ERA, Supelco | 0227-14-08, LC06435 | ERA 0227-14-08: 60 µl & Supelco LC06435: 60 µl | 0.1122 | TEBUTHIURON | E | 0.028 | 0.194 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.7291 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-09-30 | ug/L | 1.4163 | 2010-07-21 06:30 | 20102030005 | Spex Organics | T1090901009 | Spex Organics T1090901009: 20 µl | 0.11227 | TEBUTHIURON | . | 0.028 | 0.159 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.4163 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-02-09 | ug/L | 1.1255 | 2011-01-13 06:30 | 20110130056 | Supelco | LB77168 | Supelco LB77168: 80 µl | 0.12972 | TEBUTHIURON | E | 0.028 | 0.146 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1255 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-06-20 | ug/L | 1.7212 | 2011-06-02 07:30 | 20111540007 | Restek | A075606 | Restek A075606: 80 µl | 0.13247 | TEBUTHIURON | E | 0.028 | 0.228 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.7212 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-09-01 | ug/L | 1.2277 | 2011-08-10 08:00 | 20112220126 | Restek | A075606 | Restek A075606: 80 µl | 0.1344 | TEBUTHIURON | E | 0.028 | 0.165 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2277 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-09-13 | ug/L | 0.72396 | 2011-08-18 08:46 | 20112310032 | Restek | A075606 | Restek A075606: 80 µl | 0.1344 | TEBUTHIURON | E | 0.028 | 0.0973 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72396 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-11-01 | ug/L | 1.4435 | 2011-10-20 07:00 | 20112930085 | Restek | A075606 | Restek A075606: 80 µl | 0.1344 | TEBUTHIURON | . | 0.028 | 0.194 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.4435 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-05-15 | ug/L | 1.3616 | 2012-05-10 06:30 | 20121310027 | Restek | A075606 | Restek A075606: 80 µl | 0.1344 | TEBUTHIURON | . | 0.028 | 0.183 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3616 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-07-14 | ug/L | 1.2574 | 2011-06-29 07:00 | 20111810004 | Supelco | LB77168 | Supelco LB77168: 80 µl | 0.1344 | TEBUTHIURON | E | 0.028 | 0.169 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2574 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-10-02 | ug/L | 1.0417 | 2011-09-15 07:30 | 20112590005 | Supelco | LB77168 | Supelco LB77168: 80 µl | 0.1344 | TEBUTHIURON | . | 0.028 | 0.14 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0417 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-11-17 | ug/L | 1.3988 | 2011-11-09 07:00 | 20113130085 | Supelco | LB77168 | Supelco LB77168: 80 µl | 0.1344 | TEBUTHIURON | E | 0.028 | 0.188 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3988 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2011-01-28 | ug/L | 1.2352 | 2010-12-09 06:00 | 20103430087 | Restek | A075606 | Restek A075606: 80 µl | 0.1352 | TEBUTHIURON | . | 0.028 | 0.167 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2352 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-11-14 | ug/L | 1.1639 | 2012-10-30 07:00 | 20123040157 | Restek, Crescent | A088380, 05BB029 | Restek A088380: 50 µl & Crescent 05BB029: 50 µl | 0.1495 | TEBUTHIURON | . | 0.028 | 0.174 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1639 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-10-18 | ug/L | 1.0957 | 2012-10-10 07:00 | 20122850005 | Supelco, Abs | LB92783, 052412 | Supelco LB92783: 50 µl & Abs 052412: 50 µl | 0.1515 | TEBUTHIURON | E | 0.028 | 0.166 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0957 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-01-17 | ug/L | 1.1221 | 2013-01-10 06:00 | 20130100066 | Supelco, Abs | LB92783, 052412 | Supelco LB92783: 50 µl & Abs 052412: 50 µl | 0.1515 | TEBUTHIURON | . | 0.028 | 0.17 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1221 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-04-22 | ug/L | 0.94389 | 2013-04-04 06:30 | 20130940097 | Abs, ERA | 052412, 0507-12-04.2 | Abs 052412: 50 µl & ERA 0507-12-04.2: 50 µl | 0.1515 | TEBUTHIURON | . | 0.028 | 0.143 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94389 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-09-14 | ug/L | 1.2079 | 2013-08-29 06:30 | 20132420001 | Restek, Abs, Crescent | A093651, 012513, 10802AL | Restek A093651: 50 µl & Abs 012513: 50 µl & Crescent 10802AL: 5 µl | 0.1515 | TEBUTHIURON | . | 0.028 | 0.183 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2079 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-01-31 | ug/L | 0.83168 | 2014-01-23 06:30 | 20140230082 | Restek, Abs, Crescent | A093651, 012513, 10802AL | Restek A093651: 50 µl & Abs 012513: 50 µl & Crescent 10802AL: 10 µl | 0.1515 | TEBUTHIURON | . | 0.028 | 0.126 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83168 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-07-10 | ug/L | 2.2253 | 2009-05-14 08:30 | 20091340060 | Supelco | LB60731 | Supelco LB60731: 25 µl | 0.15594 | TEBUTHIURON | . | 0.02 | 0.347 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 2 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-05-06 | ug/L | 0.97636 | 2009-04-09 08:30 | 20090990155 | Restek | A063025 | Restek A063025: 25 µl | 0.1567 | TEBUTHIURON | . | 0.02 | 0.153 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97636 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-01-27 | ug/L | 1.8747 | 2009-09-17 07:00 | 20092610019 | Supelco | LB60731 | Supelco LB60731: 25 µl | 0.15683 | TEBUTHIURON | . | 0.028 | 0.294 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.8747 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-09-29 | ug/L | 1.29 | 2009-07-16 07:30 | 20091970075 | Supelco | LB60731 | Supelco LB60731: 25 µl | 0.15736 | TEBUTHIURON | . | 0.02 | 0.203 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.29 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-12-11 | ug/L | 1.3314 | 2009-08-13 07:30 | 20092260012 | Restek | A063025 | Restek A063025: 25 µl | 0.15848 | TEBUTHIURON | . | 0.028 | 0.211 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3314 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-02-11 | ug/L | 1.0996 | 2009-01-29 06:00 | 20090300003 | Restek | A063025 | Restek A063025: 25 µl | 0.1637 | TEBUTHIURON | . | 0.02 | 0.18 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0996 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-03-16 | ug/L | 1.6674 | 2009-03-05 06:30 | 20090640123 | Supelco | LB60731 | Supelco LB60731: 25 µl | 0.16973 | TEBUTHIURON | . | 0.02 | 0.283 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.6674 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-08-19 | ug/L | 1.5779 | 2013-08-01 07:00 | 20132140002 | Supelco, Crescent | LB99392, 04CC101 | Supelco LB99392: 50 µl & Crescent 04CC101: 50 µl | 0.199 | TEBUTHIURON | E | 0.028 | 0.314 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5779 | Y | Estimated Values | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-09-24 | ug/L | 1.4065 | 2013-09-12 08:30 | 20132560001 | ERA, Abs, Crescent | 0318-13-02.2, 042413, 10802AL | ERA 0318-13-02.2: 50 µl & Abs 042413: 50 µl & Crescent 10802AL: 10 µl | 0.2005 | TEBUTHIURON | . | 0.028 | 0.282 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.4065 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-10-19 | ug/L | 1.5062 | 2013-09-25 08:00 | 20132690010 | Abs, Spex Organics, Crescent | 042413, T1130514013, 10802AL | Abs 042413: 50 µl & Spex Organics T1130514013: 50 µl & Crescent 10802AL: 20 µl | 0.2005 | TEBUTHIURON | . | 0.028 | 0.302 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.5062 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-02-17 | ug/L | 1.0391 | 2009-11-05 08:30 | 20093100007 | Ultra | CF-3558 | Ultra CF-3558: 40 µl | 0.21654 | TEBUTHIURON | . | 0.028 | 0.225 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0391 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-08-10 | ug/L | 1.3784 | 2010-06-10 06:30 | 20101650033 | Ultra | CF-3558 | Ultra CF-3558: 40 µl | 0.2249 | TEBUTHIURON | . | 0.028 | 0.31 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3784 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-12-22 | ug/L | 0.95993 | 2010-09-16 06:30 | 20102630169 | Spex Organics | T1090901009 | Spex Organics T1090901009: 40 µl | 0.22606 | TEBUTHIURON | . | 0.028 | 0.217 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95993 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2009-08-25 | ug/L | 1.3207 | 2009-06-18 08:00 | 20091700144 | Restek | A063025 | Restek A063025: 40 µl | 0.25593 | TEBUTHIURON | . | 0.02 | 0.338 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3207 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2013-02-06 | ug/L | 0.90301 | 2013-01-24 06:00 | 20130250061 | Restek, Crescent | A088380, 05BB029 | Restek A088380: 100 µl & Crescent 05BB029: 100 µl | 0.299 | TEBUTHIURON | . | 0.028 | 0.27 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90301 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-12-20 | ug/L | 1.087 | 2012-12-06 06:30 | 20123420074 | Spex Organics, Crescent | T1120530013, 05BB029 | Spex Organics T1120530013: 100 µl & Crescent 05BB029: 100 µl | 0.299 | TEBUTHIURON | . | 0.028 | 0.325 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.087 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2012-12-06 | ug/L | 1.099 | 2012-11-20 06:30 | 20123260037 | Abs, ERA | 052412, 0507-12-04.2 | Abs 052412: 100 µl & ERA 0507-12-04.2: 100 µl | 0.303 | TEBUTHIURON | . | 0.028 | 0.333 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.099 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2014-01-18 | ug/L | 1.2718 | 2014-01-09 06:30 | 20140100085 | Abs, Spex Organics, Crescent | 042413, T1130514013, 10802AL | Abs 042413: 100 µl & Spex Organics T1130514013: 100 µl & Crescent 10802AL: 20 µl | 0.401 | TEBUTHIURON | . | 0.028 | 0.51 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2718 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046
TEBUTHIURON | 2033 | 82670GCM35 | 2010-07-15 | ug/L | 1.2852 | 2010-05-13 06:30 | 20101340154 | Spex Organics | T1090901009 | Spex Organics T1090901009: 100 µl | 0.55556 | TEBUTHIURON | . | 0.028 | 0.714 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2852 | Y | Recoverable Result | Y |  | 97 | 1.3488 | 0.33406 | 1.29 | 0.53416 | 0.39596 | 1.3046