## OBSP Data
AnalyteName | Schedulecode | Parm+Method Code | Analysis Date | ReportUnit | RECOVERY | Login Date | Laboratory ID | ManufShortName | CALotNum | SpikeMix | Expected Conc. | newanalyte | Remark Code | REPORTING LEVEL | Lab Value | UNITS | All | Comments | newlevel | Pct Recovery | newspike | manuf1 | manuf2 | manuf3 | lotnum1 | lotnum2 | lotnum3 | lotnum_1 | lotnum_2 | lotnum_3 | Sample Spiked | STATUS | PLOT | RANK | G | H | I | J | perm_e | N | Mean | Std-Dev. | Median | IQR | F_Pseudo | maxout
BROMACIL | 4433 | 30234GCM99 | 2014-08-20 | ug/L | 0.77882 | 2014-07-17 06:00 | 20141990081 | Abs, ERA | 032814, 0227-14-07.1 | Abs 032814: 30 µl & ERA 0227-14-07.1: 30 µl | 0.321 | BROMACIL | . | 0.16 | 0.25 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77882 | Ab032814, ER02207.1 | Ab | ER | ER | 032814 | 02207.1 | 02207.1 |  | 02207.1 | 02207.1 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.80218 | 0.033042 | 0.80218 | 0.046729 | 0.03464
BROMACIL | 4433 | 30234GCM99 | 2014-10-20 | ug/L | 0.82555 | 2014-09-25 06:30 | 20142720001 | Abs, ERA | 032814, 0227-14-07.1 | Abs 032814: 30 µl & ERA 0227-14-07.1: 30 µl | 0.321 | BROMACIL | . | 0.16 | 0.265 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82555 | Ab032814, ER02207.1 | Ab | ER | ER | 032814 | 02207.1 | 02207.1 |  | 02207.1 | 02207.1 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.80218 | 0.033042 | 0.80218 | 0.046729 | 0.03464
BROMACIL | 4433 | 30234GCM99 | 2014-12-03 | ug/L | 1.1168 | 2014-11-13 07:00 | 20143180046 | Abs, Supelco | 032814, LC06436 | Abs 032814: 60 µl & Supelco LC06436: 60 µl | 0.6402 | BROMACIL | . | 0.16 | 0.715 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1168 | Ab032814, SuLC06436 | Ab | Su | Su | 032814 | LC06436 | LC06436 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 1.1168 |  | 1.1168 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2010-05-12 | ug/L | 0.78843 | 2010-04-01 08:00 | 20100920047 | Accu | 209081217 | Accu 209081217: 40 µl | 2.3211 | BROMACIL | E | 0.8 | 1.83 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78843 | Accu 209081217 | [9 empty] | Y | Estimated Values | Y | 2 | [5 empty] | 2 | 0.77981 | 0.012186 | 0.77981 | 0.017233 | 0.012775
BROMACIL | 4433 | 30234GCM99 | 2010-08-21 | ug/L | 0.7712 | 2010-07-13 08:30 | 20101950138 | Accu | 209081217 | Accu 209081217: 80 µl | 4.318 | BROMACIL | E | 0.8 | 3.33 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7712 | Accu 209081217 | [9 empty] | Y | Estimated Values | Y | 2 | [5 empty] | 2 | 0.77981 | 0.012186 | 0.77981 | 0.017233 | 0.012775
BROMACIL | 4433 | 30234GCM99 | 2013-08-06 | ug/L | 0.68923 | 2013-06-13 07:30 | 20131650075 | ChemService, Abs | 1466200, 050613WWUF | ChemService 1466200: 50 µl & Abs 050613WWUF: 50 µl | 2.452 | BROMACIL | . | 0.16 | 1.69 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.68923 | Ch1466200, Ab050WWUF | Ch | Ab | Ab | 1466200 | 050WWUF | 050WWUF |  | 050WWUF | 050WWUF | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.68923 |  | 0.68923 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-10-22 | ug/L | 0.71633 | 2013-09-25 07:30 | 20132690008 | ChemService, Supelco | 1466200, LC00152 | ChemService 1466200: 20 µl & Supelco LC00152: 20 µl | 0.49 | BROMACIL | . | 0.16 | 0.351 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.71633 | Ch1466200, SuLC00152 | Ch | Su | Su | 1466200 | LC00152 | LC00152 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.72551 | 0.012988 | 0.72551 | 0.018367 | 0.013616
BROMACIL | 4433 | 30234GCM99 | 2014-04-23 | ug/L | 0.73469 | 2014-01-30 08:00 | 20140300033 | ChemService, Supelco | 1466200, LC00152 | ChemService 1466200: 20 µl & Supelco LC00152: 20 µl | 0.49 | BROMACIL | . | 0.16 | 0.36 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.73469 | Ch1466200, SuLC00152 | Ch | Su | Su | 1466200 | LC00152 | LC00152 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.72551 | 0.012988 | 0.72551 | 0.018367 | 0.013616
BROMACIL | 4433 | 30234GCM99 | 2015-09-29 | ug/L | 1.3134 | 2015-07-16 06:30 | 20151970168 | ChemService, Abs | 3982800, 041615 | ChemService 3982800: 20 µl & Abs 041615: 20 µl | 0.2406 | BROMACIL | . | 0.16 | 0.316 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.3134 | Ch3982800, Ab041615 | Ch | Ab | Ab | 3982800 | 041615 | 041615 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.1315 | 0.25716 | 1.1315 | 0.36367 | 0.26959
BROMACIL | 4433 | 30234GCM99 | 2016-01-26 | ug/L | 0.94971 | 2015-11-24 08:00 | 20153280156 | ChemService, Abs | 3982800, 041615 | ChemService 3982800: 40 µl & Abs 041615: 40 µl | 0.4812 | BROMACIL | . | 0.16 | 0.457 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94971 | Ch3982800, Ab041615 | Ch | Ab | Ab | 3982800 | 041615 | 041615 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.1315 | 0.25716 | 1.1315 | 0.36367 | 0.26959
BROMACIL | 4433 | 30234GCM99 | 2015-10-02 | ug/L | 0.94763 | 2015-08-27 09:00 | 20152390223 | ChemService, Crescent | 3982800, 04EE006 | ChemService 3982800: 20 µl & Crescent 04EE006: 20 µl | 0.2406 | BROMACIL | . | 0.16 | 0.228 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94763 | Ch3982800, Cr04EE006 | Ch | Cr | Cr | 3982800 | 04EE006 | 04EE006 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.94763 |  | 0.94763 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-02-06 | ug/L | 0.7713 | 2013-01-10 07:00 | 20130100068 | ChemService, Crescent | 426100, 05BB045 | ChemService 426100: 80 µl & Crescent 05BB045: 80 µl | 2.0096 | BROMACIL | . | 0.16 | 1.55 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7713 | Ch426100, Cr05BB045 | Ch | Cr | Cr | 426100 | 05BB045 | 05BB045 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.7713 |  | 0.7713 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-05-30 | ug/L | 0.94646 | 2013-04-11 07:00 | 20131020082 | ChemService, Spex Organics | 426100, T1120521007 | ChemService 426100: 40 µl & Spex Organics T1120521007: 40 µl | 1.0048 | BROMACIL | . | 0.16 | 0.951 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94646 | Ch426100, SpT111007 | Ch | Sp | Sp | 426100 | T111007 | T111007 |  | T111007 | T111007 | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.94646 |  | 0.94646 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2014-09-17 | ug/L | 0.98131 | 2014-08-18 06:45 | 20142340162 | Crescent, ERA | 03DD005, 0227-14-07.1 | Crescent 03DD005: 30 µl & ERA 0227-14-07.1: 30 µl | 0.321 | BROMACIL | . | 0.16 | 0.315 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.98131 | Cr03DD005, ER02207.1 | Cr | ER | ER | 03DD005 | 02207.1 | 02207.1 |  | 02207.1 | 02207.1 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.0569 | 0.10684 | 1.0569 | 0.15109 | 0.112
BROMACIL | 4433 | 30234GCM99 | 2015-02-26 | ug/L | 1.1324 | 2015-01-29 07:00 | 20150290077 | Crescent, ERA | 03DD005, 0227-14-07.1 | Crescent 03DD005: 60 µl & ERA 0227-14-07.1: 60 µl | 0.642 | BROMACIL | . | 0.16 | 0.727 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1324 | Cr03DD005, ER02207.1 | Cr | ER | ER | 03DD005 | 02207.1 | 02207.1 |  | 02207.1 | 02207.1 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.0569 | 0.10684 | 1.0569 | 0.15109 | 0.112
BROMACIL | 4433 | 30234GCM99 | 2014-07-16 | ug/L | 0.78569 | 2014-06-19 07:00 | 20141710103 | Crescent, Supelco | 03DD005, LC06436 | Crescent 03DD005: 60 µl & Supelco LC06436: 60 µl | 0.6402 | BROMACIL | . | 0.16 | 0.503 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78569 | Cr03DD005, SuLC06436 | Cr | Su | Su | 03DD005 | LC06436 | LC06436 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.78569 |  | 0.78569 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-12-04 | ug/L | 0.95355 | 2013-11-21 07:00 | 20133260003 | Crescent, Abs | 04CC095, 050613WWUF | Crescent 04CC095: 20 µl & Abs 050613WWUF: 20 µl | 0.4908 | BROMACIL | . | 0.16 | 0.468 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95355 | Cr04CC095, Ab050WWUF | Cr | Ab | Ab | 04CC095 | 050WWUF | 050WWUF |  | 050WWUF | 050WWUF | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.89548 | 0.082121 | 0.89548 | 0.11614 | 0.086091
BROMACIL | 4433 | 30234GCM99 | 2014-04-23 | ug/L | 0.83741 | 2014-03-27 07:00 | 20140860103 | Crescent, Abs | 04CC095, 050613WWUF | Crescent 04CC095: 20 µl & Abs 050613WWUF: 20 µl | 0.4908 | BROMACIL | . | 0.16 | 0.411 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83741 | Cr04CC095, Ab050WWUF | Cr | Ab | Ab | 04CC095 | 050WWUF | 050WWUF |  | 050WWUF | 050WWUF | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.89548 | 0.082121 | 0.89548 | 0.11614 | 0.086091
BROMACIL | 4433 | 30234GCM99 | 2012-08-24 | ug/L | 0.896 | 2012-08-07 07:00 | 20122210008 | Crescent, Restek | 05BB045, A088824 | Crescent 05BB045: 20 µl & Restek A088824: 20 µl | 0.5 | BROMACIL | . | 0.16 | 0.448 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.896 | Cr05BB045, ReA088824 | Cr | Re | Re | 05BB045 | A088824 | A088824 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.959 | 0.089095 | 0.959 | 0.126 | 0.093403
BROMACIL | 4433 | 30234GCM99 | 2012-12-18 | ug/L | 1.022 | 2012-10-30 08:00 | 20123040158 | Crescent, Restek | 05BB045, A088824 | Crescent 05BB045: 20 µl & Restek A088824: 20 µl | 0.5 | BROMACIL | . | 0.16 | 0.511 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.022 | Cr05BB045, ReA088824 | Cr | Re | Re | 05BB045 | A088824 | A088824 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.959 | 0.089095 | 0.959 | 0.126 | 0.093403
BROMACIL | 4433 | 30234GCM99 | 2010-12-15 | ug/L | 0.78196 | 2010-10-28 09:00 | 20103020152 | Crescent | 06MM136 | Crescent 06MM136: 50 µl | 2.4298 | BROMACIL | . | 0.16 | 1.9 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78196 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2011-04-14 | ug/L | 1.0831 | 2011-03-10 07:30 | 20110690080 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.97864 | BROMACIL | . | 0.16 | 1.06 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0831 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2011-07-07 | ug/L | 0.79098 | 2011-06-23 08:30 | 20111750002 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.976 | BROMACIL | . | 0.16 | 0.772 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79098 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2011-08-17 | ug/L | 0.94262 | 2011-08-04 09:30 | 20112170170 | Crescent | 06MM136 | Crescent 06MM136: 80 µl | 3.904 | BROMACIL | . | 0.16 | 3.68 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94262 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2011-11-22 | ug/L | 0.88627 | 2011-09-01 09:00 | 20112450021 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.976 | BROMACIL | . | 0.16 | 0.865 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88627 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2012-01-16 | ug/L | 1.0476 | 2011-11-03 07:30 | 20113080004 | Crescent | 06MM136 | Crescent 06MM136: 80 µl | 3.904 | BROMACIL | . | 0.16 | 4.09 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0476 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2012-03-06 | ug/L | 0.87295 | 2012-01-19 08:00 | 20120200071 | Crescent | 06MM136 | Crescent 06MM136: 50 µl | 2.44 | BROMACIL | . | 0.16 | 2.13 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.87295 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2012-03-07 | ug/L | 1.0451 | 2012-02-16 09:00 | 20120470077 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.976 | BROMACIL | . | 0.16 | 1.02 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0451 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2012-04-10 | ug/L | 0.90164 | 2012-03-08 07:30 | 20120690028 | Crescent | 06MM136 | Crescent 06MM136: 50 µl | 2.44 | BROMACIL | . | 0.16 | 2.2 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90164 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2012-07-27 | ug/L | 0.97592 | 2012-05-17 07:30 | 20121390007 | Crescent | 06MM136 | Crescent 06MM136: 80 µl | 3.904 | BROMACIL | . | 0.16 | 3.81 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97592 | Crescent 06MM136 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 10 | 0.93282 | 0.10538 | 0.92213 | 0.17213 | 0.1276
BROMACIL | 4433 | 30234GCM99 | 2010-08-18 | ug/L | 0.8179 | 2010-05-20 08:00 | 20101410010 | Crescent | 08LL076 | Crescent 08LL076: 40 µl | 2.1029 | BROMACIL | . | 0.8 | 1.72 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8179 | Crescent 08LL076 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.81423 | 0.0051912 | 0.81423 | 0.0073414 | 0.0054421
BROMACIL | 4433 | 30234GCM99 | 2010-09-29 | ug/L | 0.81056 | 2010-08-26 08:00 | 20102390261 | Crescent | 08LL076 | Crescent 08LL076: 40 µl | 2.1467 | BROMACIL | . | 0.8 | 1.74 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81056 | Crescent 08LL076 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.81423 | 0.0051912 | 0.81423 | 0.0073414 | 0.0054421
BROMACIL | 4433 | 30234GCM99 | 2014-07-17 | ug/L | 0.67913 | 2014-06-26 06:30 | 20141770150 | ERA, ChemService | 0227-14-07.1, 2616300 | ERA 0227-14-07.1: 30 µl & ChemService 2616300: 30 µl | 0.321 | BROMACIL | . | 0.16 | 0.218 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.67913 | ER02207.1, Ch2616300 | ER | Ch | Ch | 02207.1 | 2616300 | 2616300 | 02207.1 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.80841 | 0.18283 | 0.80841 | 0.25857 | 0.19167
BROMACIL | 4433 | 30234GCM99 | 2014-12-04 | ug/L | 0.93769 | 2014-11-26 07:00 | 20143300047 | ERA, ChemService | 0227-14-07.1, 2616300 | ERA 0227-14-07.1: 60 µl & ChemService 2616300: 60 µl | 0.642 | BROMACIL | . | 0.16 | 0.602 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93769 | ER02207.1, Ch2616300 | ER | Ch | Ch | 02207.1 | 2616300 | 2616300 | 02207.1 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.80841 | 0.18283 | 0.80841 | 0.25857 | 0.19167
BROMACIL | 4433 | 30234GCM99 | 2014-08-21 | ug/L | 0.96262 | 2014-07-31 06:00 | 20142130131 | ERA, Spex Organics | 0227-14-07.1, E1140402006 | ERA 0227-14-07.1: 30 µl & Spex Organics E1140402006: 30 µl | 0.321 | BROMACIL | . | 0.16 | 0.309 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96262 | ER02207.1, SpE112006 | ER | Sp | Sp | 02207.1 | E112006 | E112006 | 02207.1 | E112006 | E112006 | Y | Recoverable Result | Y | 2 | [5 empty] | 3 | 0.87747 | 0.090789 | 0.88785 | 0.18069 | 0.13394
BROMACIL | 4433 | 30234GCM99 | 2014-09-23 | ug/L | 0.78193 | 2014-09-03 06:30 | 20142470040 | ERA, Spex Organics | 0227-14-07.1, E1140402006 | ERA 0227-14-07.1: 60 µl & Spex Organics E1140402006: 60 µl | 0.642 | BROMACIL | . | 0.16 | 0.502 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78193 | ER02207.1, SpE112006 | ER | Sp | Sp | 02207.1 | E112006 | E112006 | 02207.1 | E112006 | E112006 | Y | Recoverable Result | Y | 2 | [5 empty] | 3 | 0.87747 | 0.090789 | 0.88785 | 0.18069 | 0.13394
BROMACIL | 4433 | 30234GCM99 | 2015-04-22 | ug/L | 0.88785 | 2015-04-09 06:00 | 20150990066 | ERA, Spex Organics | 0227-14-07.1, E1140402006 | ERA 0227-14-07.1: 30 µl & Spex Organics E1140402006: 30 µl | 0.321 | BROMACIL | . | 0.16 | 0.285 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88785 | ER02207.1, SpE112006 | ER | Sp | Sp | 02207.1 | E112006 | E112006 | 02207.1 | E112006 | E112006 | Y | Recoverable Result | Y | 2 | [5 empty] | 3 | 0.87747 | 0.090789 | 0.88785 | 0.18069 | 0.13394
BROMACIL | 4433 | 30234GCM99 | 2015-09-30 | ug/L | 1.1596 | 2015-07-30 09:00 | 20152110199 | ERA, ChemService | 0406-15-02, 3982800 | ERA 0406-15-02: 40 µl & ChemService 3982800: 40 µl | 0.4812 | BROMACIL | . | 0.16 | 0.558 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1596 | ER0405-02, Ch3982800 | ER | Ch | Ch | 0405-02 | 3982800 | 3982800 | 0405-02 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.01 | 0.2116 | 1.01 | 0.29925 | 0.22183
BROMACIL | 4433 | 30234GCM99 | 2016-01-12 | ug/L | 0.86035 | 2015-11-10 08:30 | 20153160002 | ERA, ChemService | 0406-15-02, 3982800 | ERA 0406-15-02: 80 µl & ChemService 3982800: 80 µl | 0.9624 | BROMACIL | . | 0.16 | 0.828 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86035 | ER0405-02, Ch3982800 | ER | Ch | Ch | 0405-02 | 3982800 | 3982800 | 0405-02 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.01 | 0.2116 | 1.01 | 0.29925 | 0.22183
BROMACIL | 4433 | 30234GCM99 | 2015-10-15 | ug/L | 0.90833 | 2015-09-10 09:00 | 20152530087 | ERA, Spex Organics | 0406-15-02, EN150529012 | ERA 0406-15-02: 20 µl & Spex Organics EN150529012: 20 µl | 0.24 | BROMACIL | . | 0.16 | 0.218 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90833 | ER0405-02, SpEN19012 | ER | Sp | Sp | 0405-02 | EN19012 | EN19012 | 0405-02 | EN19012 | EN19012 | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.90833 |  | 0.90833 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-09-26 | ug/L | 0.89242 | 2013-09-12 08:00 | 20132560002 | NSI, Abs | 042213WWUF, 050613WWUF | NSI 042213WWUF: 20 µl & Abs 050613WWUF: 20 µl | 0.4908 | BROMACIL | . | 0.16 | 0.438 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.89242 | NS042WWUF, Ab050WWUF | NS | Ab | Ab | 042WWUF | 050WWUF | 050WWUF | 042WWUF | 050WWUF | 050WWUF | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.89242 |  | 0.89242 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-06-12 | ug/L | 0.84898 | 2013-05-09 07:30 | 20131290118 | NSI, ChemService | 042213WWUF, 1466200 | NSI 042213WWUF: 50 µl & ChemService 1466200: 50 µl | 1.225 | BROMACIL | . | 0.16 | 1.04 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84898 | NS042WWUF, Ch1466200 | NS | Ch | Ch | 042WWUF | 1466200 | 1466200 | 042WWUF |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.77857 | 0.099572 | 0.77857 | 0.14082 | 0.10439
BROMACIL | 4433 | 30234GCM99 | 2014-02-05 | ug/L | 0.70816 | 2014-01-16 08:00 | 20140160088 | NSI, ChemService | 042213WWUF, 1466200 | NSI 042213WWUF: 20 µl & ChemService 1466200: 20 µl | 0.49 | BROMACIL | . | 0.16 | 0.347 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70816 | NS042WWUF, Ch1466200 | NS | Ch | Ch | 042WWUF | 1466200 | 1466200 | 042WWUF |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.77857 | 0.099572 | 0.77857 | 0.14082 | 0.10439
BROMACIL | 4433 | 30234GCM99 | 2012-12-12 | ug/L | 0.79021 | 2012-09-27 06:30 | 20122720049 | NSI, ChemService | 061412-01, 426100 | NSI 061412-01: 40 µl & ChemService 426100: 40 µl | 1.0048 | BROMACIL | . | 0.16 | 0.794 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79021 | NS0612-01, Ch426100 | NS | Ch | Ch | 0612-01 | 426100 | 426100 | 0612-01 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.79021 |  | 0.79021 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2012-12-20 | ug/L | 0.655 | 2012-12-06 07:00 | 20123420075 | NSI, Restek | 061412-01, A088824 | NSI 061412-01: 40 µl & Restek A088824: 40 µl | 1 | BROMACIL | . | 0.16 | 0.655 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.655 | NS0612-01, ReA088824 | NS | Re | Re | 0612-01 | A088824 | A088824 | 0612-01 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.655 |  | 0.655 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2010-07-15 | ug/L | 0.7222 | 2010-05-13 10:30 | 20101340153 | NWQL | BNA2SP094D | NWQL BNA2SP094D: 60 µl | 0.95541 | BROMACIL | E | 0.8 | 0.69 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7222 | NWQL BNA2SP094D | [9 empty] | Y | Estimated Values | Y | 2 | [5 empty] | 2 | 0.81219 | 0.12727 | 0.81219 | 0.17998 | 0.13342
BROMACIL | 4433 | 30234GCM99 | 2010-08-20 | ug/L | 0.90218 | 2010-07-08 10:30 | 20101900001 | NWQL | BNA2SP094D | NWQL BNA2SP094D: 60 µl | 0.97541 | BROMACIL | E | 0.8 | 0.88 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90218 | NWQL BNA2SP094D | [9 empty] | Y | Estimated Values | Y | 2 | [5 empty] | 2 | 0.81219 | 0.12727 | 0.81219 | 0.17998 | 0.13342
BROMACIL | 4433 | 30234GCM99 | 2011-05-19 | ug/L | 1.1703 | 2011-05-05 08:30 | 20111260151 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 50 µl | 0.8058 | BROMACIL | . | 0.16 | 0.943 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1703 | NWQL BNA2SP124B | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.0695 | 0.14249 | 1.0695 | 0.20151 | 0.14938
BROMACIL | 4433 | 30234GCM99 | 2011-10-18 | ug/L | 0.96875 | 2011-08-25 11:10 | 20112380056 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 80 µl | 1.28 | BROMACIL | . | 0.16 | 1.24 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96875 | NWQL BNA2SP124B | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.0695 | 0.14249 | 1.0695 | 0.20151 | 0.14938
BROMACIL | 4433 | 30234GCM99 | 2015-07-22 | ug/L | 1.0688 | 2015-06-11 06:00 | 20151620180 | Spex Organics, Abs | EN150529012, 041615 | Spex Organics EN150529012: 40 µl & Abs 041615: 40 µl | 0.48 | BROMACIL | . | 0.16 | 0.513 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0688 | SpEN19012, Ab041615 | Sp | Ab | Ab | EN19012 | 041615 | 041615 | EN19012 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 1.0688 |  | 1.0688 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2015-11-17 | ug/L | 1.05 | 2015-10-22 08:30 | 20152950079 | Spex Organics, Crescent | EN150529012, 04EE006 | Spex Organics EN150529012: 20 µl & Crescent 04EE006: 20 µl | 0.24 | BROMACIL | . | 0.16 | 0.252 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.05 | SpEN19012, Cr04EE006 | Sp | Cr | Cr | EN19012 | 04EE006 | 04EE006 | EN19012 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 1.05 |  | 1.05 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2012-07-31 | ug/L | 0.98726 | 2012-07-12 08:00 | 20121950066 | Spex Organics, ChemService | T1120521007, 426100 | Spex Organics T1120521007: 20 µl & ChemService 426100: 20 µl | 0.5024 | BROMACIL | . | 0.16 | 0.496 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.98726 | SpT111007, Ch426100 | Sp | Ch | Ch | T111007 | 426100 | 426100 | T111007 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.82305 | 0.23223 | 0.82305 | 0.32842 | 0.24346
BROMACIL | 4433 | 30234GCM99 | 2013-03-27 | ug/L | 0.65884 | 2013-02-14 08:00 | 20130460007 | Spex Organics, ChemService | T1120521007, 426100 | Spex Organics T1120521007: 20 µl & ChemService 426100: 20 µl | 0.5024 | BROMACIL | . | 0.16 | 0.331 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.65884 | SpT111007, Ch426100 | Sp | Ch | Ch | T111007 | 426100 | 426100 | T111007 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.82305 | 0.23223 | 0.82305 | 0.32842 | 0.24346
BROMACIL | 4433 | 30234GCM99 | 2013-08-01 | ug/L | 0.49307 | 2013-07-11 07:00 | 20131920148 | Spex Organics, Abs | T1130419010, 050613WWUF | Spex Organics T1130419010: 100 µl & Abs 050613WWUF: 100 µl | 2.454 | BROMACIL | . | 0.16 | 1.21 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.49307 | SpT119010, Ab050WWUF | Sp | Ab | Ab | T119010 | 050WWUF | 050WWUF | T119010 | 050WWUF | 050WWUF | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.62653 | 0.18873 | 0.62653 | 0.26691 | 0.19786
BROMACIL | 4433 | 30234GCM99 | 2014-04-22 | ug/L | 0.75998 | 2014-03-06 08:00 | 20140650080 | Spex Organics, Abs | T1130419010, 050613WWUF | Spex Organics T1130419010: 20 µl & Abs 050613WWUF: 20 µl | 0.4908 | BROMACIL | . | 0.16 | 0.373 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75998 | SpT119010, Ab050WWUF | Sp | Ab | Ab | T119010 | 050WWUF | 050WWUF | T119010 | 050WWUF | 050WWUF | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.62653 | 0.18873 | 0.62653 | 0.26691 | 0.19786
BROMACIL | 4433 | 30234GCM99 | 2013-11-14 | ug/L | 0.82041 | 2013-11-07 07:30 | 20133110109 | Spex Organics, ChemService | T1130419010, 1466200 | Spex Organics T1130419010: 20 µl & ChemService 1466200: 20 µl | 0.49 | BROMACIL | . | 0.16 | 0.402 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82041 | SpT119010, Ch1466200 | Sp | Ch | Ch | T119010 | 1466200 | 1466200 | T119010 |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.82041 |  | 0.82041 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2013-03-28 | ug/L | 0.715 | 2013-02-28 07:30 | 20130590064 | Supelco, Restek | LB92782, A088824 | Supelco LB92782: 80 µl & Restek A088824: 80 µl | 2 | BROMACIL | . | 0.16 | 1.43 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.715 | SuLB92782, ReA088824 | Su | Re | Re | LB92782 | A088824 | A088824 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.715 |  | 0.715 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2014-05-22 | ug/L | 0.75914 | 2014-05-01 06:30 | 20141220026 | Supelco, Abs | LC06436, 032814 | Supelco LC06436: 30 µl & Abs 032814: 30 µl | 0.3201 | BROMACIL | . | 0.16 | 0.243 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75914 | SuLC06436, Ab032814 | Su | Ab | Ab | LC06436 | 032814 | 032814 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 0.75914 |  | 0.75914 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2014-07-09 | ug/L | 0.81537 | 2014-06-12 06:30 | 20141640077 | Supelco, ChemService | LC06436, 2616300 | Supelco LC06436: 30 µl & ChemService 2616300: 30 µl | 0.3201 | BROMACIL | . | 0.16 | 0.261 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81537 | SuLC06436, Ch2616300 | Su | Ch | Ch | LC06436 | 2616300 | 2616300 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.0575 | 0.3424 | 1.0575 | 0.48422 | 0.35895
BROMACIL | 4433 | 30234GCM99 | 2015-04-21 | ug/L | 1.2996 | 2015-03-26 06:00 | 20150850162 | Supelco, ChemService | LC06436, 2616300 | Supelco LC06436: 30 µl & ChemService 2616300: 30 µl | 0.3201 | BROMACIL | . | 0.16 | 0.416 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2996 | SuLC06436, Ch2616300 | Su | Ch | Ch | LC06436 | 2616300 | 2616300 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.0575 | 0.3424 | 1.0575 | 0.48422 | 0.35895
BROMACIL | 4433 | 30234GCM99 | 2014-04-24 | ug/L | 0.7154 | 2014-04-10 06:00 | 20141010007 | Supelco, Crescent | LC06436, 03DD005 | Supelco LC06436: 30 µl & Crescent 03DD005: 30 µl | 0.3201 | BROMACIL | E | 0.16 | 0.229 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7154 | SuLC06436, Cr03DD005 | Su | Cr | Cr | LC06436 | 03DD005 | 03DD005 |  |  |  | Y | Estimated Values | Y | 2 | [5 empty] | 1 | 0.7154 |  | 0.7154 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2014-06-04 | ug/L | 0.47329 | 2014-05-16 06:30 | 20141360095 | Supelco, Spex Organics | LC06436, E1140402006 | Supelco LC06436: 60 µl & Spex Organics E1140402006: 60 µl | 0.6402 | BROMACIL | E | 0.16 | 0.303 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.47329 | SuLC06436, SpE112006 | Su | Sp | Sp | LC06436 | E112006 | E112006 |  | E112006 | E112006 | Y | Estimated Values | Y | 2 | [5 empty] | 2 | 0.63808 | 0.23305 | 0.63808 | 0.32958 | 0.24432
BROMACIL | 4433 | 30234GCM99 | 2014-10-21 | ug/L | 0.80287 | 2014-10-02 07:00 | 20142760025 | Supelco, Spex Organics | LC06436, E1140402006 | Supelco LC06436: 60 µl & Spex Organics E1140402006: 60 µl | 0.6402 | BROMACIL | . | 0.16 | 0.514 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80287 | SuLC06436, SpE112006 | Su | Sp | Sp | LC06436 | E112006 | E112006 |  | E112006 | E112006 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 0.63808 | 0.23305 | 0.63808 | 0.32958 | 0.24432
BROMACIL | 4433 | 30234GCM99 | 2015-11-05 | ug/L | 1.0495 | 2015-10-08 08:00 | 20152810121 | Supelco, ChemService | LC13546, 3982800 | Supelco LC13546: 80 µl & ChemService 3982800: 80 µl | 0.9624 | BROMACIL | . | 0.16 | 1.01 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0495 | SuLC13546, Ch3982800 | Su | Ch | Ch | LC13546 | 3982800 | 3982800 |  |  |  | Y | Recoverable Result | Y | 2 | [5 empty] | 1 | 1.0495 |  | 1.0495 | 0 | 0
BROMACIL | 4433 | 30234GCM99 | 2015-09-29 | ug/L | 1.2167 | 2015-06-25 06:30 | 20151760207 | Supelco, Spex Organics | LC13546, EN150529012 | Supelco LC13546: 20 µl & Spex Organics EN150529012: 20 µl | 0.24 | BROMACIL | . | 0.16 | 0.292 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2167 | SuLC13546, SpEN19012 | Su | Sp | Sp | LC13546 | EN19012 | EN19012 |  | EN19012 | EN19012 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.1146 | 0.14437 | 1.1146 | 0.20417 | 0.15135
BROMACIL | 4433 | 30234GCM99 | 2016-01-14 | ug/L | 1.0125 | 2015-12-17 08:00 | 20153510110 | Supelco, Spex Organics | LC13546, EN150529012 | Supelco LC13546: 40 µl & Spex Organics EN150529012: 40 µl | 0.48 | BROMACIL | . | 0.16 | 0.486 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0125 | SuLC13546, SpEN19012 | Su | Sp | Sp | LC13546 | EN19012 | EN19012 |  | EN19012 | EN19012 | Y | Recoverable Result | Y | 2 | [5 empty] | 2 | 1.1146 | 0.14437 | 1.1146 | 0.20417 | 0.15135
BROMACIL | 4433 | 30234GCM99 | 2011-02-03 | ug/L | 0.86799 | 2010-11-30 08:30 | 20103350004 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.98733 | BROMACIL | . | 0.16 | 0.857 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86799 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2011-04-26 | ug/L | 0.75982 | 2011-03-31 08:30 | 20110910083 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.9897 | BROMACIL | . | 0.16 | 0.752 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75982 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2011-06-29 | ug/L | 0.74582 | 2011-06-02 09:00 | 20111540005 | Supelco | LB77171 | Supelco LB77171: 80 µl | 3.9286 | BROMACIL | . | 0.16 | 2.93 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.74582 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2011-08-07 | ug/L | 0.80051 | 2011-07-14 09:30 | 20111960036 | Supelco | LB77171 | Supelco LB77171: 80 µl | 3.96 | BROMACIL | . | 0.16 | 3.17 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80051 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2011-10-04 | ug/L | 0.98384 | 2011-08-18 08:50 | 20112310029 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.99 | BROMACIL | . | 0.16 | 0.974 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.98384 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2011-11-22 | ug/L | 0.73131 | 2011-09-15 08:30 | 20112590002 | Supelco | LB77171 | Supelco LB77171: 50 µl | 2.475 | BROMACIL | . | 0.16 | 1.81 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.73131 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2012-03-01 | ug/L | 0.81313 | 2011-12-15 08:30 | 20113500022 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.99 | BROMACIL | . | 0.16 | 0.805 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81313 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042
BROMACIL | 4433 | 30234GCM99 | 2012-07-26 | ug/L | 0.77172 | 2012-05-03 07:00 | 20121250043 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.99 | BROMACIL | . | 0.16 | 0.764 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77172 | Supelco LB77171 | [9 empty] | Y | Recoverable Result | Y | 2 | [5 empty] | 8 | 0.80927 | 0.082784 | 0.78611 | 0.087742 | 0.065042