## OBSP Data
AnalyteName | Schedulecode | Parm+Method Code | Analysis Date | ReportUnit | RECOVERY | Login Date | Laboratory ID | ManufShortName | CALotNum | SpikeMix | Expected Conc. | newanalyte | Remark Code | REPORTING LEVEL | Lab Value | UNITS | All | Comments | Spike Level | Pct Recovery | Sample Spiked | STATUS | PLOT | G | H | I | J | perm_e | N | Mean | Std-Dev. | Median | IQR | F_Pseudo | Median Pct Recovery | maxout
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-03-24 | ug/L |  | 2010-01-21 07:00 | 20100210100 | Accu | 209081187 | Accu 209081187: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 98 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-05-07 | ug/L |  | 2010-02-25 06:30 | 20100570086 | Accu | 209081187 | Accu 209081187: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-09-09 | ug/L |  | 2010-06-24 06:30 | 20101760006 | Accu | 209081187 | Accu 209081187: 40 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-12-14 | ug/L |  | 2010-09-02 06:30 | 20102500051 | Accu | 209081187 | Accu 209081187: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-06-26 | ug/L |  | 2010-04-22 06:30 | 20101130055 | ChemService | 430-36C | ChemService 430-36C: 40 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-09-21 | ug/L |  | 2010-07-01 06:30 | 20101830110 | ChemService | 430-36C | ChemService 430-36C: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-06-03 | ug/L |  | 2010-04-01 06:30 | 20100920048 | Crescent | 08Ll064 | Crescent 08Ll064: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-07-21 | ug/L |  | 2010-05-20 06:30 | 20101410006 | Crescent | 08Ll064 | Crescent 08Ll064: 40 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-10-19 | ug/L |  | 2010-08-03 06:30 | 20102160002 | Crescent | 08Ll064 | Crescent 08Ll064: 20 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-09-13 | ug/L |  | 2011-08-18 08:47 | 20112310030 | ERA | 0318-13-02.2 | ERA 0318-13-02.2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-10-12 | ug/L |  | 2012-09-27 07:00 | 20122720048 | ERA | 0507-12-04.2 | ERA 0507-12-04.2: 50 µl |  | CHLORPYRIFOS | < | 0.01 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-13 | ug/L |  | 2010-11-04 07:30 | 20103080116 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-08 | ug/L |  | 2011-01-06 07:00 | 20110070109 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-28 | ug/L |  | 2011-02-10 07:00 | 20110410061 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-04-13 | ug/L |  | 2011-03-31 07:30 | 20110910084 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-06-20 | ug/L |  | 2011-05-26 07:30 | 20111470065 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-07-28 | ug/L |  | 2011-07-14 08:00 | 20111960039 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-10-25 | ug/L |  | 2011-10-13 07:30 | 20112870073 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-12-03 | ug/L |  | 2011-11-17 07:00 | 20113220198 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-03-05 | ug/L |  | 2012-02-16 07:30 | 20120470078 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-05-08 | ug/L |  | 2012-04-26 07:00 | 20121180062 | ERA | 0818-10-05 LS 2033#2 | ERA 0818-10-05 LS 2033#2: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-03-19 | ug/L |  | 2010-01-14 06:30 | 20100150090 | Restek | A069639 | Restek A069639: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-07-29 | ug/L |  | 2010-05-27 06:30 | 20101480005 | Restek | A069639 | Restek A069639: 100 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-10-18 | ug/L |  | 2010-08-05 06:30 | 20102180092 | Restek | A069639 | Restek A069639: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-09 | ug/L |  | 2010-10-21 07:30 | 20102950072 | Restek | A075606 | Restek A075606: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-28 | ug/L |  | 2010-12-09 06:00 | 20103430087 | Restek | A075606 | Restek A075606: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-03-08 | ug/L |  | 2011-02-24 08:00 | 20110560005 | Restek | A075606 | Restek A075606: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-04-23 | ug/L |  | 2011-04-14 07:30 | 20111050150 | Restek | A075606 | Restek A075606: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-06-20 | ug/L |  | 2011-06-02 07:30 | 20111540007 | Restek | A075606 | Restek A075606: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-08-17 | ug/L |  | 2011-07-21 07:30 | 20112030057 | Restek | A075606 | Restek A075606: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-09-01 | ug/L |  | 2011-08-10 08:00 | 20112220126 | Restek | A075606 | Restek A075606: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-09-13 | ug/L |  | 2011-08-18 08:46 | 20112310032 | Restek | A075606 | Restek A075606: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-11-01 | ug/L |  | 2011-10-20 07:00 | 20112930085 | Restek | A075606 | Restek A075606: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-12-04 | ug/L |  | 2011-11-22 07:30 | 20113260111 | Restek | A075606 | Restek A075606: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-02-14 | ug/L |  | 2012-02-02 12:10 | 20120370001 | Restek | A075606 | Restek A075606: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-05-15 | ug/L |  | 2012-05-10 06:30 | 20121310027 | Restek | A075606 | Restek A075606: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-06-28 | ug/L |  | 2012-06-21 06:00 | 20121730123 | Restek | A088380 | Restek A088380: 50 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-06-16 | ug/L |  | 2010-04-08 06:30 | 20100990014 | Spex Organics | T1090901009 | Spex Organics T1090901009: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-07-15 | ug/L |  | 2010-05-13 06:30 | 20101340154 | Spex Organics | T1090901009 | Spex Organics T1090901009: 100 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-09-30 | ug/L |  | 2010-07-21 06:30 | 20102030005 | Spex Organics | T1090901009 | Spex Organics T1090901009: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-12-22 | ug/L |  | 2010-09-16 06:30 | 20102630169 | Spex Organics | T1090901009 | Spex Organics T1090901009: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-11 | ug/L |  | 2010-10-14 07:30 | 20102880001 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-21 | ug/L |  | 2010-11-30 07:00 | 20103350003 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-19 | ug/L |  | 2011-01-27 06:30 | 20110280004 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-03-23 | ug/L |  | 2011-03-10 06:30 | 20110690078 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-05-14 | ug/L |  | 2011-05-05 07:30 | 20111260154 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-07-07 | ug/L |  | 2011-06-23 07:00 | 20111750006 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-09-19 | ug/L |  | 2011-08-25 10:00 | 20112380055 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-09-29 | ug/L |  | 2011-09-08 07:30 | 20112520209 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-11-17 | ug/L |  | 2011-11-03 07:00 | 20113080001 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-12-21 | ug/L |  | 2011-12-13 07:30 | 20113490056 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-01-27 | ug/L |  | 2012-01-19 07:00 | 20120200073 | Spex Organics | T1100628017 | Spex Organics T1100628017: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-04-05 | ug/L |  | 2012-03-22 06:30 | 20120830077 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-06-02 | ug/L |  | 2012-05-24 07:00 | 20121460001 | Spex Organics | T1100628017 | Spex Organics T1100628017: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-18 | ug/L |  | 2010-11-10 07:30 | 20103140135 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-09 | ug/L |  | 2011-01-13 06:30 | 20110130056 | Supelco | LB77168 | Supelco LB77168: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-03-30 | ug/L |  | 2011-03-17 08:00 | 20110770001 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-05-25 | ug/L |  | 2011-05-12 07:00 | 20111330148 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-07-14 | ug/L |  | 2011-06-29 07:00 | 20111810004 | Supelco | LB77168 | Supelco LB77168: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-07-20 | ug/L |  | 2011-07-07 06:30 | 20111890005 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-09-24 | ug/L |  | 2011-09-01 07:30 | 20112450022 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-10-02 | ug/L |  | 2011-09-15 07:30 | 20112590005 | Supelco | LB77168 | Supelco LB77168: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-11-17 | ug/L |  | 2011-11-09 07:00 | 20113130085 | Supelco | LB77168 | Supelco LB77168: 80 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-12-22 | ug/L |  | 2011-12-15 07:30 | 20113500021 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-03-09 | ug/L |  | 2012-03-01 07:30 | 20120620015 | Supelco | LB77168 | Supelco LB77168: 40 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-05-19 | ug/L |  | 2010-03-11 06:30 | 20100710178 | Ultra | CF-3558 | Ultra CF-3558: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-08-10 | ug/L |  | 2010-06-10 06:30 | 20101650033 | Ultra | CF-3558 | Ultra CF-3558: 40 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-12-01 | ug/L |  | 2010-08-26 06:30 | 20102390088 | Ultra | CF-3558 | Ultra CF-3558: 20 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-05-12 | ug/L |  | 2015-04-30 07:00 | 20151210037 | Abs, Abs | 041415, 040615 | Abs 041415: 20 µl & Abs 040615: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-07-21 | ug/L |  | 2015-07-09 07:00 | 20151900132 | Abs, Abs | 040615, 041415 | Abs 040615: 40 µl & Abs 041415: 40 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-09-05 | ug/L |  | 2015-08-20 06:00 | 20152330039 | Spex Organics, Abs | EN150513008, 041415 | Spex Organics EN150513008: 40 µl & Abs 041415: 40 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-06-11 | ug/L |  | 2013-05-30 07:30 | 20131510026 | ERA, Crescent | 0318-13-02.2, 04CC101 | ERA 0318-13-02.2: 20 µl & Crescent 04CC101: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-08-19 | ug/L |  | 2013-08-01 07:00 | 20132140002 | Supelco, Crescent | LB99392, 04CC101 | Supelco LB99392: 50 µl & Crescent 04CC101: 50 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-06-26 | ug/L |  | 2014-05-29 06:00 | 20141490205 | Abs, ERA | 032714, 0227-14-08 | Abs 032714: 30 µl & ERA 0227-14-08: 30 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-05-28 | ug/L |  | 2014-05-16 06:00 | 20141360094 | Spex Organics, ERA | E1140407002, 0227-14-08 | Spex Organics E1140407002: 30 µl & ERA 0227-14-08: 30 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-07-09 | ug/L |  | 2014-06-26 06:00 | 20141770149 | Spex Organics, ERA | E1140407002, 0227-14-08 | Spex Organics E1140407002: 60 µl & ERA 0227-14-08: 60 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-09-05 | ug/L |  | 2012-08-16 06:30 | 20122300021 | Spex Organics, ERA | T1120530013, 0507-12-04.2 | Spex Organics T1120530013: 20 µl & ERA 0507-12-04.2: 20 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-01-15 | ug/L |  | 2012-12-20 06:30 | 20123560092 | Spex Organics, ERA | T1120530013, 0507-12-04.2 | Spex Organics T1120530013: 50 µl & ERA 0507-12-04.2: 50 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-10-21 | ug/L |  | 2014-10-09 06:00 | 20142820151 | Abs, NSI | 032714, 042814 | Abs 032714: 30 µl & NSI 042814: 30 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-12-04 | ug/L |  | 2014-11-19 06:00 | 20143240095 | Abs, NSI | 032714, 042814 | Abs 032714: 60 µl & NSI 042814: 60 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-09-22 | ug/L |  | 2014-09-03 06:00 | 20142470041 | Spex Organics, NSI | E1140407002, 042814 | Spex Organics E1140407002: 30 µl & NSI 042814: 30 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-02-21 | ug/L |  | 2015-02-12 06:00 | 20150430092 | Spex Organics, NSI | E1140407002, 042814 | Spex Organics E1140407002: 60 µl & NSI 042814: 60 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2016-01-21 | ug/L |  | 2015-10-01 06:00 | 20152750006 | Abs, Spex Organics | 040615, EN150514012 | Abs 040615: 80 µl & Spex Organics EN150514012: 80 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2016-02-02 | ug/L |  | 2015-11-10 06:00 | 20153160001 | Spex Organics, Spex Organics | EN150513008, EN150514012 | Spex Organics EN150513008: 20 µl & Spex Organics EN150514012: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-07-14 | ug/L |  | 2012-07-03 08:00 | 20121870008 | Restek, Supelco | A088380, LB92783 | Restek A088380: 20 µl & Supelco LB92783: 20 µl |  | CHLORPYRIFOS | < | 0.0036 | 0.0036 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-11-20 | ug/L |  | 2012-11-08 06:30 | 20123140095 | Spex Organics, Supelco | T1120530013, LB92783 | Spex Organics T1120530013: 100 µl & Supelco LB92783: 100 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-02-22 | ug/L |  | 2013-02-07 06:30 | 20130390105 | Spex Organics, Supelco | T1120530013, LB92783 | Spex Organics T1120530013: 20 µl & Supelco LB92783: 20 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-04-02 | ug/L |  | 2013-03-21 06:30 | 20130810108 | Spex Organics, Supelco | T1120530013, LB92783 | Spex Organics T1120530013: 50 µl & Supelco LB92783: 50 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-09-24 | ug/L |  | 2013-09-12 08:30 | 20132560001 | ERA, Abs, Crescent | 0318-13-02.2, 042413, 10802AL | ERA 0318-13-02.2: 50 µl & Abs 042413: 50 µl & Crescent 10802AL: 10 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-01-10 | ug/L |  | 2013-12-23 06:30 | 20133580036 | ERA, Abs, Crescent | 0318-13-02.2, 042413, 10802AL | ERA 0318-13-02.2: 20 µl & Abs 042413: 20 µl & Crescent 10802AL: 10 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-11-23 | ug/L |  | 2013-11-14 07:30 | 20133190012 | Supelco, Abs, Crescent | LB99392, 042413, 10802AL | Supelco LB99392: 20 µl & Abs 042413: 20 µl & Crescent 10802AL: 10 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-03-05 | ug/L |  | 2014-02-20 06:30 | 20140510043 | Supelco, Crescent, Crescent | LB99392, 04CC101, 10802AL | Supelco LB99392: 20 µl & Crescent 04CC101: 20 µl & Crescent 10802AL: 10 µl |  | CHLORPYRIFOS | < | 0.01 | 0.01 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2016-01-28 | ug/L | 0.6 | 2015-10-22 06:00 | 20152950078 | Abs, ChemService | 041415, 3851600 | Abs 041415: 20 µl & ChemService 3851600: 20 µl | 0.014 | CHLORPYRIFOS | . | 0.01 | 0.0084 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.6 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-07-10 | ug/L | 1.0214 | 2015-06-25 06:00 | 20151760208 | Spex Organics, ChemService | EN150514012, 3851600 | Spex Organics EN150514012: 20 µl & ChemService 3851600: 20 µl | 0.014 | CHLORPYRIFOS | . | 0.01 | 0.0143 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0214 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2016-01-26 | ug/L | 0.86429 | 2015-10-29 06:00 | 20153030068 | ChemService, Spex Organics | 3851600, EN150513008 | ChemService 3851600: 20 µl & Spex Organics EN150513008: 20 µl | 0.014 | CHLORPYRIFOS | . | 0.01 | 0.0121 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86429 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-07-31 | ug/L | 0.77896 | 2015-07-23 06:00 | 20152040177 | Abs, ERA | 041415, 0406-15-01 | Abs 041415: 20 µl & ERA 0406-15-01: 20 µl | 0.01502 | CHLORPYRIFOS | . | 0.01 | 0.0117 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77896 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-08-19 | ug/L | 0.60586 | 2015-08-06 06:00 | 20152190003 | ERA, Spex Organics | 0406-15-01, EN150513008 | ERA 0406-15-01: 20 µl & Spex Organics EN150513008: 20 µl | 0.01502 | CHLORPYRIFOS | . | 0.01 | 0.0091 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.60586 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-04-29 | ug/L | 0.83582 | 2014-04-17 06:00 | 20141080004 | Crescent, Abs | 03DD006, 032714 | Crescent 03DD006: 30 µl & Abs 032714: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0168 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83582 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-01-31 | ug/L | 0.75622 | 2015-01-15 06:00 | 20150150115 | Supelco, Abs | LC06435, 032714 | Supelco LC06435: 30 µl & Abs 032714: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0152 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75622 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-06-24 | ug/L | 1.1244 | 2014-05-22 07:00 | 20141420140 | ERA, Crescent | 0227-14-08, 03DD006 | ERA 0227-14-08: 30 µl & Crescent 03DD006: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0226 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1244 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-08-16 | ug/L | 0.74129 | 2014-08-07 07:00 | 20142200164 | NSI, Crescent | 042814, 03DD006 | NSI 042814: 30 µl & Crescent 03DD006: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0149 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.74129 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-06-30 | ug/L | 0.88557 | 2014-06-12 06:00 | 20141640076 | Crescent, Spex Organics | 03DD006, E1140407002 | Crescent 03DD006: 30 µl & Spex Organics E1140407002: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0178 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88557 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-07-17 | ug/L | 0.67662 | 2014-07-10 06:00 | 20141910243 | Supelco, Spex Organics | LC06435, E1140407002 | Supelco LC06435: 30 µl & Spex Organics E1140407002: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0136 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.67662 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-07-16 | ug/L | 0.70647 | 2014-07-02 07:00 | 20141840097 | ERA, Supelco | 0227-14-08, LC06435 | ERA 0227-14-08: 30 µl & Supelco LC06435: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0142 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70647 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-12-15 | ug/L | 0.47761 | 2014-12-03 06:00 | 20143370094 | NSI, Supelco | 042814, LC06435 | NSI 042814: 30 µl & Supelco LC06435: 30 µl | 0.0201 | CHLORPYRIFOS | . | 0.01 | 0.0096 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.47761 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-03-05 | ug/L | 0.92929 | 2011-02-17 08:30 | 20110490035 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 25 µl | 0.02475 | CHLORPYRIFOS | . | 0.0036 | 0.023 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.92929 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-06-17 | ug/L | 0.95714 | 2015-06-04 07:00 | 20151550293 | ChemService, Abs | 3851600, 040615 | ChemService 3851600: 40 µl & Abs 040615: 40 µl | 0.028 | CHLORPYRIFOS | . | 0.01 | 0.0268 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95714 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-05-12 | ug/L | 0.86786 | 2015-05-07 07:00 | 20151270112 | Abs, ChemService | 041415, 3851600 | Abs 041415: 40 µl & ChemService 3851600: 40 µl | 0.028 | CHLORPYRIFOS | . | 0.01 | 0.0243 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86786 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2016-02-08 | ug/L | 0.64913 | 2015-12-09 06:00 | 20153440001 | Spex Organics, ERA | EN150514012, 0406-15-01 | Spex Organics EN150514012: 40 µl & ERA 0406-15-01: 40 µl | 0.03004 | CHLORPYRIFOS | . | 0.01 | 0.0195 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.64913 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2016-02-09 | ug/L | 0.66911 | 2015-12-17 06:00 | 20153510109 | ERA, Spex Organics | 0406-15-01, EN150513008 | ERA 0406-15-01: 40 µl & Spex Organics EN150513008: 40 µl | 0.03004 | CHLORPYRIFOS | . | 0.01 | 0.0201 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.66911 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-06-26 | ug/L | 0.95 | 2013-06-13 06:30 | 20131650073 | Abs, ChemService | 042413, 1473000 | Abs 042413: 20 µl & ChemService 1473000: 20 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0361 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-07-11 | ug/L | 0.95 | 2013-06-27 06:30 | 20131790009 | ChemService, ERA | 1473000, 0318-13-02.2 | ChemService 1473000: 20 µl & ERA 0318-13-02.2: 20 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0361 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-09-05 | ug/L | 1.0158 | 2013-08-15 06:30 | 20132280016 | Crescent, Spex Organics | 04CC101, T1130514013 | Crescent 04CC101: 20 µl & Spex Organics T1130514013: 20 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0386 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0158 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-01-08 | ug/L | 0.95526 | 2013-12-12 06:30 | 20133470008 | ChemService, Supelco | 1473000, LB99392 | ChemService 1473000: 20 µl & Supelco LB99392: 20 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0363 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95526 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-11-09 | ug/L | 1.0842 | 2013-10-31 06:30 | 20133050050 | Spex Organics, Supelco | T1130514013, LB99392 | Spex Organics T1130514013: 20 µl & Supelco LB99392: 20 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0412 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0842 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-12-05 | ug/L | 0.85789 | 2013-11-26 07:00 | 20133310030 | Crescent, ChemService, Cresc | 04CC101, 1473000, 10802AL | Crescent 04CC101: 20 µl & ChemService 1473000: 20 µl & Crescent 10802AL: 20 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0326 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.85789 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-04-03 | ug/L | 0.86053 | 2014-03-13 07:00 | 20140730138 | Crescent, ChemService, Cresc | 04CC101, 1473000, 10802AL | Crescent 04CC101: 20 µl & ChemService 1473000: 20 µl & Crescent 10802AL: 10 µl | 0.038 | CHLORPYRIFOS | . | 0.01 | 0.0327 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86053 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-06-30 | ug/L | 0.78386 | 2010-04-29 10:30 | 20101190091 | NWQL | PSP2SP116C | NWQL PSP2SP116C: 40 µl | 0.038782 | CHLORPYRIFOS | . | 0.01 | 0.0304 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78386 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-08-27 | ug/L | 0.64179 | 2014-08-14 06:00 | 20142270001 | Crescent, Abs | 03DD006, 032714 | Crescent 03DD006: 60 µl & Abs 032714: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0258 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.64179 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-05-20 | ug/L | 1.0224 | 2014-05-01 06:00 | 20141220025 | Supelco, Abs | LC06435, 032714 | Supelco LC06435: 60 µl & Abs 032714: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0411 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0224 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-03-31 | ug/L | 0.86567 | 2015-03-12 06:00 | 20150720069 | ERA, Crescent | 0227-14-08, 03DD006 | ERA 0227-14-08: 60 µl & Crescent 03DD006: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0348 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86567 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-10-30 | ug/L | 0.70896 | 2014-10-23 06:00 | 20142960086 | NSI, Crescent | 042814, 03DD006 | NSI 042814: 60 µl & Crescent 03DD006: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0285 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70896 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-11-19 | ug/L | 0.77861 | 2014-11-06 06:00 | 20143110108 | Spex Organics, Crescent | E1140407002, 03DD006 | Spex Organics E1140407002: 60 µl & Crescent 03DD006: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0313 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77861 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-10-04 | ug/L | 0.72388 | 2014-09-25 06:00 | 20142720002 | Supelco, Spex Organics | LC06435, E1140407002 | Supelco LC06435: 60 µl & Spex Organics E1140407002: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0291 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72388 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-06-28 | ug/L | 0.88806 | 2014-06-05 07:00 | 20141570088 | ERA, Supelco | 0227-14-08, LC06435 | ERA 0227-14-08: 60 µl & Supelco LC06435: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0357 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88806 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-09-23 | ug/L | 0.60199 | 2014-09-10 06:00 | 20142530087 | NSI, Supelco | 042814, LC06435 | NSI 042814: 60 µl & Supelco LC06435: 60 µl | 0.0402 | CHLORPYRIFOS | . | 0.01 | 0.0242 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.60199 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-12-16 | ug/L | 0.93251 | 2010-09-09 06:40 | 20102530008 | NWQL | PSP2SP116C | NWQL PSP2SP116C: 40 µl | 0.040428 | CHLORPYRIFOS | . | 0.0036 | 0.0377 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93251 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-08-25 | ug/L | 1.0482 | 2010-06-17 10:30 | 20101720149 | NWQL | PSP2SP116C | NWQL PSP2SP116C: 40 µl | 0.040547 | CHLORPYRIFOS | . | 0.01 | 0.0425 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0482 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-12 | ug/L | 0.77762 | 2010-10-28 07:30 | 20103020155 | Accu | 210061341 | Accu 210061341: 40 µl | 0.046552 | CHLORPYRIFOS | . | 0.0036 | 0.0362 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77762 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-14 | ug/L | 0.7105 | 2011-01-20 07:30 | 20110210095 | Accu | 210061341 | Accu 210061341: 40 µl | 0.047432 | CHLORPYRIFOS | . | 0.0036 | 0.0337 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7105 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-07-02 | ug/L | 0.64639 | 2011-06-16 08:00 | 20111680035 | Accu | 210061341 | Accu 210061341: 40 µl | 0.048113 | CHLORPYRIFOS | . | 0.0036 | 0.0311 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.64639 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-07-20 | ug/L | 0.87552 | 2011-07-07 07:00 | 20111890004 | Accu | 210061341 | Accu 210061341: 40 µl | 0.0482 | CHLORPYRIFOS | . | 0.0036 | 0.0422 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.87552 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-11-05 | ug/L | 0.78216 | 2011-10-27 07:30 | 20113010047 | Accu | 210061341 | Accu 210061341: 40 µl | 0.0482 | CHLORPYRIFOS | . | 0.0036 | 0.0377 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78216 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-03-27 | ug/L | 0.6805 | 2012-03-15 08:00 | 20120760111 | Accu | 210061341 | Accu 210061341: 40 µl | 0.0482 | CHLORPYRIFOS | . | 0.0036 | 0.0328 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.6805 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-04-07 | ug/L | 0.80334 | 2011-03-24 06:30 | 20110840108 | Accu | 210061341 | Accu 210061341: 40 µl | 0.048423 | CHLORPYRIFOS | . | 0.0036 | 0.0389 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80334 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-07 | ug/L | 0.56738 | 2010-10-07 07:30 | 20102810019 | Abs | 080610 | Abs 080610: 40 µl | 0.048997 | CHLORPYRIFOS | . | 0.0036 | 0.0278 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.56738 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-22 | ug/L | 0.78448 | 2011-02-02 06:30 | 20110340007 | Abs | 080610 | Abs 080610: 40 µl | 0.049332 | CHLORPYRIFOS | . | 0.0036 | 0.0387 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78448 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-05-03 | ug/L | 0.9018 | 2011-04-22 06:20 | 20111170107 | Abs | 080610 | Abs 080610: 40 µl | 0.049567 | CHLORPYRIFOS | . | 0.0036 | 0.0447 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9018 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-08-22 | ug/L | 0.80658 | 2011-07-28 07:00 | 20112100007 | Abs | 080610 | Abs 080610: 40 µl | 0.04984 | CHLORPYRIFOS | E | 0.0036 | 0.0402 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80658 | Y | Estimated Values | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-10-20 | ug/L | 0.76645 | 2011-10-07 06:30 | 20112840031 | Abs | 080610 | Abs 080610: 40 µl | 0.04984 | CHLORPYRIFOS | . | 0.0036 | 0.0382 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.76645 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-01-20 | ug/L | 0.84872 | 2012-01-12 07:00 | 20120120071 | Abs | 080610 | Abs 080610: 40 µl | 0.04984 | CHLORPYRIFOS | . | 0.0036 | 0.0423 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84872 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-03-08 | ug/L | 0.77849 | 2012-02-23 07:30 | 20120540036 | Abs | 080610 | Abs 080610: 40 µl | 0.04984 | CHLORPYRIFOS | . | 0.0036 | 0.0388 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77849 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-04-12 | ug/L | 0.70425 | 2012-03-29 08:00 | 20120930003 | Abs | 080610 | Abs 080610: 40 µl | 0.04984 | CHLORPYRIFOS | . | 0.0036 | 0.0351 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70425 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-09-24 | ug/L | 0.749 | 2015-09-17 06:00 | 20152610011 | ERA, Abs | 0406-15-01, 040615 | ERA 0406-15-01: 80 µl & Abs 040615: 80 µl | 0.06008 | CHLORPYRIFOS | . | 0.01 | 0.045 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.749 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2015-09-17 | ug/L | 0.68908 | 2015-09-03 06:00 | 20152460215 | Spex Organics, ERA | EN150514012, 0406-15-01 | Spex Organics EN150514012: 80 µl & ERA 0406-15-01: 80 µl | 0.06008 | CHLORPYRIFOS | . | 0.01 | 0.0414 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.68908 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-06-22 | ug/L | 0.79087 | 2011-06-07 10:00 | 20111590005 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 75 µl | 0.074096 | CHLORPYRIFOS | . | 0.0036 | 0.0586 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79087 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-02-12 | ug/L | 0.88737 | 2014-02-06 06:30 | 20140370077 | Spex Organics, ERA | T1130514013, 0318-13-02.2 | Spex Organics T1130514013: 50 µl & ERA 0318-13-02.2: 50 µl | 0.095 | CHLORPYRIFOS | . | 0.01 | 0.0843 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88737 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-06-05 | ug/L | 1.0347 | 2013-05-16 07:00 | 20131370001 | ChemService, Supelco | 1473000, LB99392 | ChemService 1473000: 50 µl & Supelco LB99392: 50 µl | 0.095 | CHLORPYRIFOS | . | 0.01 | 0.0983 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0347 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-10-19 | ug/L | 1.2316 | 2013-09-25 08:00 | 20132690010 | Abs, Spex Organics, Crescent | 042413, T1130514013, 10802AL | Abs 042413: 50 µl & Spex Organics T1130514013: 50 µl & Crescent 10802AL: 20 µl | 0.095 | CHLORPYRIFOS | . | 0.01 | 0.117 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.2316 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-04-22 | ug/L | 0.57807 | 2010-02-11 06:30 | 20100420055 | Supelco | LB69639 | Supelco LB69639: 20 µl | 0.09549 | CHLORPYRIFOS | . | 0.01 | 0.0552 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.57807 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-28 | ug/L | 0.67563 | 2010-11-24 07:00 | 20103330007 | Abs | 080610 | Abs 080610: 80 µl | 0.095911 | CHLORPYRIFOS | . | 0.0036 | 0.0648 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.67563 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-02-01 | ug/L | 0.7714 | 2010-12-16 06:30 | 20103510060 | Accu | 210061341 | Accu 210061341: 80 µl | 0.09593 | CHLORPYRIFOS | . | 0.0036 | 0.074 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7714 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-05-21 | ug/L | 0.93413 | 2011-04-28 07:00 | 20111190184 | Accu | 210061341 | Accu 210061341: 80 µl | 0.096025 | CHLORPYRIFOS | . | 0.0036 | 0.0897 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93413 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-03-25 | ug/L | 0.59296 | 2010-01-28 06:30 | 20100280077 | NSI | 090109-01Pest | NSI 090109-01Pest: 20 µl | 0.096127 | CHLORPYRIFOS | . | 0.01 | 0.057 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.59296 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-08-30 | ug/L | 0.82676 | 2011-08-04 08:00 | 20112170174 | Accu | 210061341 | Accu 210061341: 80 µl | 0.0964 | CHLORPYRIFOS | . | 0.0036 | 0.0797 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82676 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-12-08 | ug/L | 0.71162 | 2011-11-29 07:00 | 20113340046 | Accu | 210061341 | Accu 210061341: 80 µl | 0.0964 | CHLORPYRIFOS | . | 0.0036 | 0.0686 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.71162 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-03-02 | ug/L | 0.83817 | 2012-02-09 07:00 | 20120400058 | Accu | 210061341 | Accu 210061341: 80 µl | 0.0964 | CHLORPYRIFOS | . | 0.0036 | 0.0808 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83817 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-03-16 | ug/L | 0.82754 | 2011-03-03 07:00 | 20110620082 | Abs | 080610 | Abs 080610: 80 µl | 0.098967 | CHLORPYRIFOS | . | 0.0036 | 0.0819 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82755 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-04-20 | ug/L | 0.93484 | 2011-04-07 09:00 | 20110980030 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 100 µl | 0.099483 | CHLORPYRIFOS | . | 0.0036 | 0.093 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93484 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-05-22 | ug/L | 0.93474 | 2013-05-02 06:00 | 20131230084 | Restek, Crescent | A088380, 05BB029 | Restek A088380: 20 µl & Crescent 05BB029: 20 µl | 0.0996 | CHLORPYRIFOS | . | 0.01 | 0.0931 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93474 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-08-29 | ug/L | 0.75402 | 2012-08-07 06:00 | 20122210006 | Spex Organics, Crescent | T1120530013, 05BB029 | Spex Organics T1120530013: 20 µl & Crescent 05BB029: 20 µl | 0.0996 | CHLORPYRIFOS | . | 0.0036 | 0.0751 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.75402 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-09-17 | ug/L | 0.8755 | 2012-08-30 06:30 | 20122430227 | Crescent, ERA | 05BB029, 0507-12-04.2 | Crescent 05BB029: 20 µl & ERA 0507-12-04.2: 20 µl | 0.0996 | CHLORPYRIFOS | . | 0.0036 | 0.0872 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8755 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-03-07 | ug/L | 0.61245 | 2013-02-19 06:30 | 20130510051 | Crescent, ERA | 05BB029, 0507-12-04.2 | Crescent 05BB029: 20 µl & ERA 0507-12-04.2: 20 µl | 0.0996 | CHLORPYRIFOS | . | 0.01 | 0.061 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.61245 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-02-09 | ug/L | 0.76846 | 2012-01-26 07:00 | 20120260028 | Abs | 080610 | Abs 080610: 80 µl | 0.09968 | CHLORPYRIFOS | . | 0.0036 | 0.0766 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.76846 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-04-24 | ug/L | 0.72632 | 2012-04-12 07:00 | 20121040012 | Abs | 080610 | Abs 080610: 80 µl | 0.09968 | CHLORPYRIFOS | . | 0.0036 | 0.0724 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72632 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-08-03 | ug/L | 0.93784 | 2010-06-03 06:30 | 20101550006 | NSI | 090109-01Pest | NSI 090109-01Pest: 20 µl | 0.10012 | CHLORPYRIFOS | . | 0.01 | 0.0939 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93784 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-06-16 | ug/L | 0.90419 | 2012-06-07 06:30 | 20121600156 | Abs | 052412 | Abs 052412: 20 µl | 0.1002 | CHLORPYRIFOS | . | 0.0036 | 0.0906 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90419 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-07-23 | ug/L | 0.78543 | 2013-07-11 06:30 | 20131920149 | Restek, Abs | A093651, 012513 | Restek A093651: 20 µl & Abs 012513: 20 µl | 0.1002 | CHLORPYRIFOS | . | 0.01 | 0.0787 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78543 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-09-26 | ug/L | 0.71058 | 2012-09-13 06:00 | 20122570122 | Restek, Abs | A088380, 052412 | Restek A088380: 20 µl & Abs 052412: 20 µl | 0.1002 | CHLORPYRIFOS | . | 0.0036 | 0.0712 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.71058 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-03-18 | ug/L | 0.66766 | 2013-03-05 06:00 | 20130650005 | Restek, Abs | A088380, 052412 | Restek A088380: 20 µl & Abs 052412: 20 µl | 0.1002 | CHLORPYRIFOS | . | 0.01 | 0.0669 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.66766 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-08-14 | ug/L | 0.90319 | 2012-07-19 06:30 | 20122020004 | Abs, Supelco | 052412, LB92783 | Abs 052412: 20 µl & Supelco LB92783: 20 µl | 0.1002 | CHLORPYRIFOS | . | 0.0036 | 0.0905 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90319 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-04-15 | ug/L | 0.68463 | 2014-03-27 06:30 | 20140860102 | Restek, Abs, Crescent | A093651, 012513, 10802AL | Restek A093651: 20 µl & Abs 012513: 20 µl & Crescent 10802AL: 10 µl | 0.1002 | CHLORPYRIFOS | . | 0.01 | 0.0686 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.68463 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-11-04 | ug/L | 0.72156 | 2013-10-17 06:30 | 20132910008 | Restek, Abs, Crescent | A093651, 052412, 10802AL | Restek A093651: 20 µl & Abs 052412: 20 µl & Crescent 10802AL: 5 µl | 0.1002 | CHLORPYRIFOS | . | 0.01 | 0.0723 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72156 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-01-08 | ug/L | 0.86627 | 2013-12-05 08:00 | 20133400002 | Restek, Abs, Crescent | A093651, 052412, 10802AL | Restek A093651: 20 µl & Abs 052412: 20 µl & Crescent 10802AL: 5 µl | 0.1002 | CHLORPYRIFOS | . | 0.01 | 0.0868 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.86627 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-05-29 | ug/L | 0.83506 | 2011-05-19 07:30 | 20111390178 | Abs | 080610 | Abs 080610: 80 µl | 0.10035 | CHLORPYRIFOS | . | 0.0036 | 0.0838 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83506 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2011-01-20 | ug/L | 1.0129 | 2010-11-18 09:00 | 20103230114 | NWQL | PSP2SP117A | NWQL PSP2SP117A: 100 µl | 0.1007 | CHLORPYRIFOS | . | 0.0036 | 0.102 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0129 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-07-07 | ug/L | 0.81908 | 2010-05-06 06:30 | 20101270038 | Supelco | LB69639 | Supelco LB69639: 40 µl | 0.18802 | CHLORPYRIFOS | . | 0.01 | 0.154 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.81908 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-09-20 | ug/L | 1.1211 | 2013-09-06 06:50 | 20132490342 | Spex Organics, ERA | T1130514013, 0318-13-02.2 | Spex Organics T1130514013: 100 µl & ERA 0318-13-02.2: 100 µl | 0.19 | CHLORPYRIFOS | . | 0.01 | 0.213 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1211 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-01-18 | ug/L | 0.91579 | 2014-01-09 06:30 | 20140100085 | Abs, Spex Organics, Crescent | 042413, T1130514013, 10802AL | Abs 042413: 100 µl & Spex Organics T1130514013: 100 µl & Crescent 10802AL: 20 µl | 0.19 | CHLORPYRIFOS | . | 0.01 | 0.174 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91579 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-09-23 | ug/L | 0.88108 | 2010-07-13 06:30 | 20101950137 | Supelco | LB69639 | Supelco LB69639: 40 µl | 0.19181 | CHLORPYRIFOS | . | 0.01 | 0.169 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88108 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-11-14 | ug/L | 0.80723 | 2012-10-30 07:00 | 20123040157 | Restek, Crescent | A088380, 05BB029 | Restek A088380: 50 µl & Crescent 05BB029: 50 µl | 0.249 | CHLORPYRIFOS | . | 0.01 | 0.201 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.80723 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-10-18 | ug/L | 0.8503 | 2012-10-10 07:00 | 20122850005 | Supelco, Abs | LB92783, 052412 | Supelco LB92783: 50 µl & Abs 052412: 50 µl | 0.2505 | CHLORPYRIFOS | . | 0.01 | 0.213 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8503 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-01-17 | ug/L | 0.97804 | 2013-01-10 06:00 | 20130100066 | Supelco, Abs | LB92783, 052412 | Supelco LB92783: 50 µl & Abs 052412: 50 µl | 0.2505 | CHLORPYRIFOS | . | 0.01 | 0.245 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97804 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-04-22 | ug/L | 0.77445 | 2013-04-04 06:30 | 20130940097 | Abs, ERA | 052412, 0507-12-04.2 | Abs 052412: 50 µl & ERA 0507-12-04.2: 50 µl | 0.2505 | CHLORPYRIFOS | . | 0.01 | 0.194 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77445 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-09-14 | ug/L | 0.91018 | 2013-08-29 06:30 | 20132420001 | Restek, Abs, Crescent | A093651, 012513, 10802AL | Restek A093651: 50 µl & Abs 012513: 50 µl & Crescent 10802AL: 5 µl | 0.2505 | CHLORPYRIFOS | . | 0.01 | 0.228 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91018 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2014-01-31 | ug/L | 0.79441 | 2014-01-23 06:30 | 20140230082 | Restek, Abs, Crescent | A093651, 012513, 10802AL | Restek A093651: 50 µl & Abs 012513: 50 µl & Crescent 10802AL: 10 µl | 0.2505 | CHLORPYRIFOS | . | 0.01 | 0.199 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.79441 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2010-10-27 | ug/L | 0.74085 | 2010-08-12 06:30 | 20102250096 | NSI | 090109-01Pest | NSI 090109-01Pest: 100 µl | 0.49403 | CHLORPYRIFOS | . | 0.0036 | 0.366 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.74085 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2013-02-06 | ug/L | 0.8755 | 2013-01-24 06:00 | 20130250061 | Restek, Crescent | A088380, 05BB029 | Restek A088380: 100 µl & Crescent 05BB029: 100 µl | 0.498 | CHLORPYRIFOS | . | 0.01 | 0.436 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.8755 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-12-20 | ug/L | 0.9257 | 2012-12-06 06:30 | 20123420074 | Spex Organics, Crescent | T1120530013, 05BB029 | Spex Organics T1120530013: 100 µl & Crescent 05BB029: 100 µl | 0.498 | CHLORPYRIFOS | . | 0.01 | 0.461 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9257 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315
CHLORPYRIFOS | 2033 | 38933GCM35 | 2012-12-06 | ug/L | 0.95409 | 2012-11-20 06:30 | 20123260037 | Abs, ERA | 052412, 0507-12-04.2 | Abs 052412: 100 µl & ERA 0507-12-04.2: 100 µl | 0.501 | CHLORPYRIFOS | . | 0.01 | 0.478 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95409 | Y | Recoverable Result | Y | [5 empty] | 98 | 0.82013 | 0.13701 | 0.81315 | 0.18862 | 0.13982 | 0.81315