## OBSP Data
AnalyteName | Schedulecode | Parm+Method Code | Analysis Date | ReportUnit | RECOVERY | Login Date | Laboratory ID | ManufShortName | CALotNum | SpikeMix | Expected Conc. | newanalyte | Remark Code | REPORTING LEVEL | Lab Value | UNITS | All | Comments | Spike Level | Pct Recovery | Sample Spiked | STATUS | PLOT | G | H | I | J | perm_e | N | Mean | Std-Dev. | Median | IQR | F_Pseudo | Median Pct Recovery | maxout
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-02-15 | ug/L |  | 2010-11-18 08:00 | 20103230118 | Abs | 090910WW | Abs 090910WW: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 77 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-04-15 | ug/L |  | 2011-03-24 07:30 | 20110840109 | Abs | 090910WW | Abs 090910WW: 50 µl |  | 1-METHYLNAPHTHALENE | . | 0.04 | 0.014 | MICROGRAMS PER LITER |  |  |  |  | N | False Positives | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-05-03 | ug/L |  | 2011-04-14 08:30 | 20111050151 | Abs | 090910WW | Abs 090910WW: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-07-07 | ug/L |  | 2011-06-16 09:30 | 20111680036 | Abs | 090910WW | Abs 090910WW: 50 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-08-16 | ug/L |  | 2011-07-28 08:30 | 20112100004 | Abs | 090910WW | Abs 090910WW: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-11-22 | ug/L |  | 2011-09-08 09:30 | 20112520211 | Abs | 090910WW | Abs 090910WW: 80 µl |  | 1-METHYLNAPHTHALENE | . | 0.04 | 0.018 | MICROGRAMS PER LITER |  |  |  |  | N | False Positives | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-12-08 | ug/L |  | 2011-10-20 07:30 | 20112930084 | Abs | 090910WW | Abs 090910WW: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-03-07 | ug/L |  | 2012-01-26 08:00 | 20120260060 | Abs | 090910WW | Abs 090910WW: 50 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-04-10 | ug/L |  | 2012-03-01 08:30 | 20120620013 | Abs | 090910WW | Abs 090910WW: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-05-03 | ug/L |  | 2012-04-05 06:30 | 20120960124 | Abs | 090910WW | Abs 090910WW: 50 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-05-12 | ug/L |  | 2010-04-01 08:00 | 20100920047 | Accu | 209081217 | Accu 209081217: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-08-21 | ug/L |  | 2010-07-13 08:30 | 20101950138 | Accu | 209081217 | Accu 209081217: 80 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-12-14 | ug/L |  | 2010-10-14 09:00 | 20102880003 | ChemService | 448-12A | ChemService 448-12A: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-02-18 | ug/L |  | 2011-02-02 07:30 | 20110340009 | ChemService | 448-12A | ChemService 448-12A: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-06-28 | ug/L |  | 2011-05-19 08:30 | 20111390173 | ChemService | 448-12A | ChemService 448-12A: 80 µl |  | 1-METHYLNAPHTHALENE | . | 0.04 | 0.014 | MICROGRAMS PER LITER |  |  |  |  | N | False Positives | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-08-06 | ug/L |  | 2011-07-07 08:30 | 20111890007 | ChemService | 448-12A | ChemService 448-12A: 50 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-08-17 | ug/L |  | 2011-08-10 09:30 | 20112220129 | ChemService | 448-12A | ChemService 448-12A: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-02-01 | ug/L |  | 2011-11-29 07:30 | 20113340047 | ChemService | 448-12A | ChemService 448-12A: 50 µl |  | 1-METHYLNAPHTHALENE | E | 0.04 | 0.007 | MICROGRAMS PER LITER |  |  |  |  | N | False Positives | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-03-06 | ug/L |  | 2012-01-12 08:00 | 20120120069 | ChemService | 448-12A | ChemService 448-12A: 80 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-03-07 | ug/L |  | 2012-02-09 07:30 | 20120400055 | ChemService | 448-12A | ChemService 448-12A: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-05-03 | ug/L |  | 2012-03-22 07:30 | 20120830078 | ChemService | 448-12A | ChemService 448-12A: 50 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-05-22 | ug/L |  | 2012-04-19 07:00 | 20121110020 | ChemService | 448-12A | ChemService 448-12A: 80 µl |  | 1-METHYLNAPHTHALENE | . | 0.04 | 0.011 | MICROGRAMS PER LITER |  |  |  |  | N | False Positives | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-05-19 | ug/L |  | 2010-04-15 07:30 | 20101060002 | ChemService | 480-95B | ChemService 480-95B: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-07-29 | ug/L |  | 2012-06-21 07:00 | 20121730124 | Crescent | 05BB045 | Crescent 05BB045: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-08-18 | ug/L |  | 2010-05-20 08:00 | 20101410010 | Crescent | 08LL076 | Crescent 08LL076: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-09-29 | ug/L |  | 2010-08-26 08:00 | 20102390261 | Crescent | 08LL076 | Crescent 08LL076: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-03-11 | ug/L |  | 2010-02-11 07:30 | 20100420053 | NSI | 083109-01WW | NSI 083109-01WW: 80 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-07-23 | ug/L |  | 2010-06-03 08:30 | 20101550010 | NSI | 083109-01WW | NSI 083109-01WW: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.2 | 0.2 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-07-28 | ug/L |  | 2012-06-07 08:00 | 20121600155 | Supelco | LB92782 | Supelco LB92782: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-05-22 | ug/L |  | 2014-05-01 06:30 | 20141220026 | Supelco, Abs | LC06436, 032814 | Supelco LC06436: 30 µl & Abs 032814: 30 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-09-29 | ug/L |  | 2015-07-16 06:30 | 20151970168 | ChemService, Abs | 3982800, 041615 | ChemService 3982800: 20 µl & Abs 041615: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2016-01-26 | ug/L |  | 2015-11-24 08:00 | 20153280156 | ChemService, Abs | 3982800, 041615 | ChemService 3982800: 40 µl & Abs 041615: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-07-22 | ug/L |  | 2015-06-11 06:00 | 20151620180 | Spex Organics, Abs | EN150529012, 041615 | Spex Organics EN150529012: 40 µl & Abs 041615: 40 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-08-06 | ug/L |  | 2013-06-13 07:30 | 20131650075 | ChemService, Abs | 1466200, 050613WWUF | ChemService 1466200: 50 µl & Abs 050613WWUF: 50 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-12-04 | ug/L |  | 2013-11-21 07:00 | 20133260003 | Crescent, Abs | 04CC095, 050613WWUF | Crescent 04CC095: 20 µl & Abs 050613WWUF: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-04-23 | ug/L |  | 2014-03-27 07:00 | 20140860103 | Crescent, Abs | 04CC095, 050613WWUF | Crescent 04CC095: 20 µl & Abs 050613WWUF: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-04-24 | ug/L |  | 2014-04-10 06:00 | 20141010007 | Supelco, Crescent | LC06436, 03DD005 | Supelco LC06436: 30 µl & Crescent 03DD005: 30 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-10-02 | ug/L |  | 2015-08-27 09:00 | 20152390223 | ChemService, Crescent | 3982800, 04EE006 | ChemService 3982800: 20 µl & Crescent 04EE006: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-11-17 | ug/L |  | 2015-10-22 08:30 | 20152950079 | Spex Organics, Crescent | EN150529012, 04EE006 | Spex Organics EN150529012: 20 µl & Crescent 04EE006: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-02-06 | ug/L |  | 2013-01-10 07:00 | 20130100068 | ChemService, Crescent | 426100, 05BB045 | ChemService 426100: 80 µl & Crescent 05BB045: 80 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-08-20 | ug/L |  | 2014-07-17 06:00 | 20141990081 | Abs, ERA | 032814, 0227-14-07.1 | Abs 032814: 30 µl & ERA 0227-14-07.1: 30 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-10-20 | ug/L |  | 2014-09-25 06:30 | 20142720001 | Abs, ERA | 032814, 0227-14-07.1 | Abs 032814: 30 µl & ERA 0227-14-07.1: 30 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-09-17 | ug/L |  | 2014-08-18 06:45 | 20142340162 | Crescent, ERA | 03DD005, 0227-14-07.1 | Crescent 03DD005: 30 µl & ERA 0227-14-07.1: 30 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-02-26 | ug/L |  | 2015-01-29 07:00 | 20150290077 | Crescent, ERA | 03DD005, 0227-14-07.1 | Crescent 03DD005: 60 µl & ERA 0227-14-07.1: 60 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-08-24 | ug/L |  | 2012-08-07 07:00 | 20122210008 | Crescent, Restek | 05BB045, A088824 | Crescent 05BB045: 20 µl & Restek A088824: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-12-18 | ug/L |  | 2012-10-30 08:00 | 20123040158 | Crescent, Restek | 05BB045, A088824 | Crescent 05BB045: 20 µl & Restek A088824: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-03-28 | ug/L |  | 2013-02-28 07:30 | 20130590064 | Supelco, Restek | LB92782, A088824 | Supelco LB92782: 80 µl & Restek A088824: 80 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-10-22 | ug/L |  | 2013-09-25 07:30 | 20132690008 | ChemService, Supelco | 1466200, LC00152 | ChemService 1466200: 20 µl & Supelco LC00152: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-04-23 | ug/L |  | 2014-01-30 08:00 | 20140300033 | ChemService, Supelco | 1466200, LC00152 | ChemService 1466200: 20 µl & Supelco LC00152: 20 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-08-15 | ug/L |  | 2013-07-25 07:30 | 20132070037 | Crescent, Supelco | 04CC095, LC00152 | Crescent 04CC095: 100 µl & Supelco LC00152: 100 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-12-03 | ug/L |  | 2014-11-13 07:00 | 20143180046 | Abs, Supelco | 032814, LC06436 | Abs 032814: 60 µl & Supelco LC06436: 60 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-07-16 | ug/L |  | 2014-06-19 07:00 | 20141710103 | Crescent, Supelco | 03DD005, LC06436 | Crescent 03DD005: 60 µl & Supelco LC06436: 60 µl |  | 1-METHYLNAPHTHALENE | < | 0.04 | 0.04 | MICROGRAMS PER LITER |  |  |  |  | N | < Reporting Level | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-08-08 08:00 | ug/L |  | 2013-08-08 08:00 | 20132210039 | Spex Organics, Supelco | T1130419010, LC00152 | Spex Organics T1130419010: 20 µl & Supelco LC00152: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | R | 0.04 |  | MICROGRAMS PER LITER |  |  |  |  | Y | Deleted Values | N | [5 empty] | 0 | [5 empty] | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-08-20 | ug/L | 0.78667 | 2010-08-12 11:30 | 20102250100 | Restek | A070024 | Restek A070024: 20 µl | 0.1411 | 1-METHYLNAPHTHALENE | E | 0.2 | 0.111 | MICROGRAMS PER LITER |  |  | < Reporting Level | 0.78667 | Y | Estimated Values | Y | [5 empty] | 1 | 0.78667 |  | 0.78667 | 0 | 0 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-11-17 | ug/L | 0.99662 | 2015-10-29 08:30 | 20153030067 | Abs, ERA | 041615, 0406-15-02 | Abs 041615: 20 µl & ERA 0406-15-02: 20 µl | 0.0592 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.059 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99662 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-10-15 | ug/L | 1.1824 | 2015-09-10 09:00 | 20152530087 | ERA, Spex Organics | 0406-15-02, EN150529012 | ERA 0406-15-02: 20 µl & Spex Organics EN150529012: 20 µl | 0.0592 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.07 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1824 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-07-09 | ug/L | 0.94684 | 2015-05-21 06:30 | 20151410210 | Supelco, Crescent | LC13546, 04EE006 | Supelco LC13546: 20 µl & Crescent 04EE006: 20 µl | 0.0602 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.057 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94684 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-09-29 | ug/L | 1.0299 | 2015-06-25 06:30 | 20151760207 | Supelco, Spex Organics | LC13546, EN150529012 | Supelco LC13546: 20 µl & Spex Organics EN150529012: 20 µl | 0.0602 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.062 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0299 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-05-08 | ug/L | 0.96346 | 2015-04-30 07:30 | 20151210038 | Abs, Supelco | 041615, LC13546 | Abs 041615: 20 µl & Supelco LC13546: 20 µl | 0.0602 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.058 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96346 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-09-26 | ug/L | 0.90541 | 2013-09-12 08:00 | 20132560002 | NSI, Abs | 042213WWUF, 050613WWUF | NSI 042213WWUF: 20 µl & Abs 050613WWUF: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.067 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90541 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-04-22 | ug/L | 0.77027 | 2014-03-06 08:00 | 20140650080 | Spex Organics, Abs | T1130419010, 050613WWUF | Spex Organics T1130419010: 20 µl & Abs 050613WWUF: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.057 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77027 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-02-05 | ug/L | 0.7973 | 2014-01-16 08:00 | 20140160088 | NSI, ChemService | 042213WWUF, 1466200 | NSI 042213WWUF: 20 µl & ChemService 1466200: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.059 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7973 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-11-14 | ug/L | 1.0405 | 2013-11-07 07:30 | 20133110109 | Spex Organics, ChemService | T1130419010, 1466200 | Spex Organics T1130419010: 20 µl & ChemService 1466200: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.077 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0405 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-11-14 | ug/L | 0.83784 | 2013-10-17 07:30 | 20132910006 | NSI, Crescent | 042213WWUF, 04CC095 | NSI 042213WWUF: 20 µl & Crescent 04CC095: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.062 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83784 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-04-23 | ug/L | 0.91892 | 2014-02-06 07:30 | 20140370076 | NSI, Crescent | 042213WWUF, 04CC095 | NSI 042213WWUF: 20 µl & Crescent 04CC095: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.068 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91892 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-04-24 | ug/L | 0.89189 | 2014-04-03 08:00 | 20140930054 | Spex Organics, Supelco | T1130419010, LC00152 | Spex Organics T1130419010: 20 µl & Supelco LC00152: 20 µl | 0.074 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.066 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.89189 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-06-05 | ug/L | 0.799 | 2014-05-08 07:30 | 20141290007 | ChemService, Abs | 2616300, 032814 | ChemService 2616300: 30 µl & Abs 032814: 30 µl | 0.0801 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.064 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.799 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-05-07 | ug/L | 0.99875 | 2015-04-23 06:00 | 20151140152 | ChemService, Abs | 2616300, 032814 | ChemService 2616300: 30 µl & Abs 032814: 30 µl | 0.0801 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.08 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99875 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-07-17 | ug/L | 0.82397 | 2014-06-26 06:30 | 20141770150 | ERA, ChemService | 0227-14-07.1, 2616300 | ERA 0227-14-07.1: 30 µl & ChemService 2616300: 30 µl | 0.0801 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.066 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82397 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-07-09 | ug/L | 0.72409 | 2014-06-12 06:30 | 20141640077 | Supelco, ChemService | LC06436, 2616300 | Supelco LC06436: 30 µl & ChemService 2616300: 30 µl | 0.0801 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.058 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72409 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-04-21 | ug/L | 0.89888 | 2015-03-26 06:00 | 20150850162 | Supelco, ChemService | LC06436, 2616300 | Supelco LC06436: 30 µl & ChemService 2616300: 30 µl | 0.0801 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.072 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.89888 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-04-20 | ug/L | 0.74906 | 2015-02-26 06:00 | 20150570009 | ChemService, Crescent | 2616300, 03DD005 | ChemService 2616300: 30 µl & Crescent 03DD005: 30 µl | 0.0801 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.06 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.74906 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-09-11 | ug/L | 1.0864 | 2014-08-14 06:30 | 20142270002 | Spex Organics, Abs | E1140402006, 032814 | Spex Organics E1140402006: 30 µl & Abs 032814: 30 µl | 0.081 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.088 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0864 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-01-27 | ug/L | 0.92593 | 2014-12-18 06:00 | 20143520107 | Spex Organics, Crescent | E1140402006, 03DD005 | Spex Organics E1140402006: 30 µl & Crescent 03DD005: 30 µl | 0.081 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.075 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.92593 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-08-21 | ug/L | 1.037 | 2014-07-31 06:00 | 20142130131 | ERA, Spex Organics | 0227-14-07.1, E1140402006 | ERA 0227-14-07.1: 30 µl & Spex Organics E1140402006: 30 µl | 0.081 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.084 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.037 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-04-22 | ug/L | 1.1111 | 2015-04-09 06:00 | 20150990066 | ERA, Spex Organics | 0227-14-07.1, E1140402006 | ERA 0227-14-07.1: 30 µl & Spex Organics E1140402006: 30 µl | 0.081 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.09 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1111 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-09-30 | ug/L | 1.0304 | 2015-07-30 09:00 | 20152110199 | ERA, ChemService | 0406-15-02, 3982800 | ERA 0406-15-02: 40 µl & ChemService 3982800: 40 µl | 0.1184 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.122 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0304 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-10-01 | ug/L | 0.94595 | 2015-08-13 09:00 | 20152250271 | Abs, ERA | 041615, 0406-15-02 | Abs 041615: 40 µl & ERA 0406-15-02: 40 µl | 0.1184 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.112 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94595 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2016-01-14 | ug/L | 0.93854 | 2015-12-17 08:00 | 20153510110 | Supelco, Spex Organics | LC13546, EN150529012 | Supelco LC13546: 40 µl & Spex Organics EN150529012: 40 µl | 0.1204 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.113 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93854 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2016-01-13 | ug/L | 0.96346 | 2015-12-03 08:30 | 20153370177 | Crescent, Supelco | 04EE006, LC13546 | Crescent 04EE006: 40 µl & Supelco LC13546: 40 µl | 0.1204 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.116 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96346 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-07-31 | ug/L | 1.06 | 2012-07-12 08:00 | 20121950066 | Spex Organics, ChemService | T1120521007, 426100 | Spex Organics T1120521007: 20 µl & ChemService 426100: 20 µl | 0.15 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.159 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.06 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-03-27 | ug/L | 0.78 | 2013-02-14 08:00 | 20130460007 | Spex Organics, ChemService | T1120521007, 426100 | Spex Organics T1120521007: 20 µl & ChemService 426100: 20 µl | 0.15 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.117 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-05-29 | ug/L | 0.87333 | 2013-03-27 07:00 | 20130870086 | NSI, Crescent | 061412-01, 05BB045 | NSI 061412-01: 20 µl & Crescent 05BB045: 20 µl | 0.15 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.131 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.87333 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-09-18 | ug/L | 0.93333 | 2012-08-23 07:30 | 20122370103 | NSI, Supelco | 061412-01, LB92782 | NSI 061412-01: 20 µl & Supelco LB92782: 20 µl | 0.15 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.14 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93333 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-09-23 | ug/L | 0.89888 | 2014-09-04 07:00 | 20142480001 | ChemService, Abs | 2616300, 032814 | ChemService 2616300: 60 µl & Abs 032814: 60 µl | 0.1602 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.144 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.89888 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-12-04 | ug/L | 0.93633 | 2014-11-26 07:00 | 20143300047 | ERA, ChemService | 0227-14-07.1, 2616300 | ERA 0227-14-07.1: 60 µl & ChemService 2616300: 60 µl | 0.1602 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.15 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93633 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-08-20 | ug/L | 0.93009 | 2014-07-10 06:30 | 20141910242 | ChemService, Crescent | 2616300, 03DD005 | ChemService 2616300: 60 µl & Crescent 03DD005: 60 µl | 0.1602 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.149 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93009 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-07-08 | ug/L | 0.7963 | 2014-05-29 06:30 | 20141490201 | Spex Organics, Abs | E1140402006, 032814 | Spex Organics E1140402006: 60 µl & Abs 032814: 60 µl | 0.162 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.129 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.7963 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-11-05 | ug/L | 0.99383 | 2014-10-23 06:30 | 20142960087 | Spex Organics, Crescent | E1140402006, 03DD005 | Spex Organics E1140402006: 60 µl & Crescent 03DD005: 60 µl | 0.162 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.161 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99383 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-09-23 | ug/L | 1.0309 | 2014-09-03 06:30 | 20142470040 | ERA, Spex Organics | 0227-14-07.1, E1140402006 | ERA 0227-14-07.1: 60 µl & Spex Organics E1140402006: 60 µl | 0.162 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.167 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0309 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-06-04 | ug/L | 0.67901 | 2014-05-16 06:30 | 20141360095 | Supelco, Spex Organics | LC06436, E1140402006 | Supelco LC06436: 60 µl & Spex Organics E1140402006: 60 µl | 0.162 | 1-METHYLNAPHTHALENE | E | 0.04 | 0.11 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.67901 | Y | Estimated Values | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-10-21 | ug/L | 0.83951 | 2014-10-02 07:00 | 20142760025 | Supelco, Spex Organics | LC06436, E1140402006 | Supelco LC06436: 60 µl & Spex Organics E1140402006: 60 µl | 0.162 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.136 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83951 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-06-12 | ug/L | 0.93514 | 2013-05-09 07:30 | 20131290118 | NSI, ChemService | 042213WWUF, 1466200 | NSI 042213WWUF: 50 µl & ChemService 1466200: 50 µl | 0.185 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.173 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93514 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2014-02-05 | ug/L | 0.93514 | 2014-01-02 06:15 | 20140030013 | Spex Organics, Crescent | T1130419010, 04CC095 | Spex Organics T1130419010: 50 µl & Crescent 04CC095: 50 µl | 0.185 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.173 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93514 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-06-14 | ug/L | 1.0216 | 2013-05-23 07:30 | 20131440104 | NSI, Supelco | 042213WWUF, LC00152 | NSI 042213WWUF: 50 µl & Supelco LC00152: 50 µl | 0.185 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.189 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0216 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-05-19 | ug/L | 1.0722 | 2011-05-05 08:30 | 20111260151 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 50 µl | 0.20145 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.216 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0722 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2016-01-12 | ug/L | 1.0473 | 2015-11-10 08:30 | 20153160002 | ERA, ChemService | 0406-15-02, 3982800 | ERA 0406-15-02: 80 µl & ChemService 3982800: 80 µl | 0.2368 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.248 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0473 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-10-16 | ug/L | 1.0557 | 2015-09-24 09:00 | 20152670134 | Crescent, ERA | 04EE006, 0406-15-02 | Crescent 04EE006: 80 µl & ERA 0406-15-02: 80 µl | 0.2368 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.25 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0557 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-07-15 | ug/L | 0.65731 | 2010-05-13 10:30 | 20101340153 | NWQL | BNA2SP094D | NWQL BNA2SP094D: 60 µl | 0.23885 | 1-METHYLNAPHTHALENE | E | 0.2 | 0.157 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.65731 | Y | Estimated Values | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2015-11-05 | ug/L | 0.94269 | 2015-10-08 08:00 | 20152810121 | Supelco, ChemService | LC13546, 3982800 | Supelco LC13546: 80 µl & ChemService 3982800: 80 µl | 0.2408 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.227 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94269 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-08-20 | ug/L | 0.77506 | 2010-07-08 10:30 | 20101900001 | NWQL | BNA2SP094D | NWQL BNA2SP094D: 60 µl | 0.24385 | 1-METHYLNAPHTHALENE | E | 0.2 | 0.189 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.77506 | Y | Estimated Values | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-08-20 | ug/L | 0.73506 | 2010-08-12 08:30 | 20102250097 | Restek | A070024 | Restek A070024: 40 µl | 0.27481 | 1-METHYLNAPHTHALENE | . | 0.2 | 0.202 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.73506 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-05-13 | ug/L | 0.63934 | 2010-03-11 07:30 | 20100710176 | Ultra | CF-3634 | Ultra CF-3634: 40 µl | 0.28467 | 1-METHYLNAPHTHALENE | E | 0.2 | 0.182 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.63934 | Y | Estimated Values | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-12-12 | ug/L | 0.82 | 2012-09-27 06:30 | 20122720049 | NSI, ChemService | 061412-01, 426100 | NSI 061412-01: 40 µl & ChemService 426100: 40 µl | 0.3 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.246 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.82 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-12-20 | ug/L | 0.83 | 2012-12-06 07:00 | 20123420075 | NSI, Restek | 061412-01, A088824 | NSI 061412-01: 40 µl & Restek A088824: 40 µl | 0.3 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.249 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-05-30 | ug/L | 0.93 | 2013-04-11 07:00 | 20131020082 | ChemService, Spex Organics | 426100, T1120521007 | ChemService 426100: 40 µl & Spex Organics T1120521007: 40 µl | 0.3 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.279 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.93 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-02-07 | ug/L | 0.78333 | 2013-01-24 06:30 | 20130250062 | Spex Organics, Supelco | T1120521007, LB92782 | Spex Organics T1120521007: 40 µl & Supelco LB92782: 40 µl | 0.3 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.235 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.78333 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-10-18 | ug/L | 1.1094 | 2011-08-25 11:10 | 20112380056 | NWQL | BNA2SP124B | NWQL BNA2SP124B: 80 µl | 0.32 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.355 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.1094 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-07-07 | ug/L | 0.72892 | 2011-06-23 08:30 | 20111750002 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.332 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.242 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.72892 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-11-22 | ug/L | 0.91265 | 2011-09-01 09:00 | 20112450021 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.332 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.303 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.91265 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-03-07 | ug/L | 0.85843 | 2012-02-16 09:00 | 20120470077 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.332 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.285 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.85843 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-04-14 | ug/L | 1.0033 | 2011-03-10 07:30 | 20110690080 | Crescent | 06MM136 | Crescent 06MM136: 20 µl | 0.3329 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.334 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0033 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-02-03 | ug/L | 0.62816 | 2010-11-30 08:30 | 20103350004 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.34068 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.214 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.62816 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-04-26 | ug/L | 0.97512 | 2011-03-31 08:30 | 20110910083 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.3415 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.333 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.97512 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-10-04 | ug/L | 0.88993 | 2011-08-18 08:50 | 20112310029 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.3416 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.304 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.88993 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-03-01 | ug/L | 0.90164 | 2011-12-15 08:30 | 20113500022 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.3416 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.308 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90164 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-07-26 | ug/L | 0.99824 | 2012-05-03 07:00 | 20121250043 | Supelco | LB77171 | Supelco LB77171: 20 µl | 0.3416 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.341 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.99824 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2013-08-01 | ug/L | 0.85676 | 2013-07-11 07:00 | 20131920148 | Spex Organics, Abs | T1130419010, 050613WWUF | Spex Organics T1130419010: 100 µl & Abs 050613WWUF: 100 µl | 0.37 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.317 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.85676 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-07-20 | ug/L | 0.59964 | 2010-05-06 08:30 | 20101270037 | Restek | A070024 | Restek A070024: 80 µl | 0.552 | 1-METHYLNAPHTHALENE | E | 0.2 | 0.331 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.59964 | Y | Estimated Values | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-08-18 | ug/L | 0.70817 | 2010-06-17 08:00 | 20101720151 | Ultra | CF-3634 | Ultra CF-3634: 80 µl | 0.58743 | 1-METHYLNAPHTHALENE | . | 0.2 | 0.416 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70817 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-12-19 | ug/L | 0.83167 | 2012-11-08 07:30 | 20123140092 | Spex Organics, Supelco | T1120521007, LB92782 | Spex Organics T1120521007: 80 µl & Supelco LB92782: 80 µl | 0.6 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.499 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.83167 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2010-12-15 | ug/L | 0.70657 | 2010-10-28 09:00 | 20103020152 | Crescent | 06MM136 | Crescent 06MM136: 50 µl | 0.82653 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.584 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.70657 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-03-06 | ug/L | 0.92289 | 2012-01-19 08:00 | 20120200071 | Crescent | 06MM136 | Crescent 06MM136: 50 µl | 0.83 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.766 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.92289 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-04-10 | ug/L | 0.90361 | 2012-03-08 07:30 | 20120690028 | Crescent | 06MM136 | Crescent 06MM136: 50 µl | 0.83 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.75 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.90361 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-11-22 | ug/L | 0.94145 | 2011-09-15 08:30 | 20112590002 | Supelco | LB77171 | Supelco LB77171: 50 µl | 0.854 | 1-METHYLNAPHTHALENE | . | 0.04 | 0.804 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.94145 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-08-17 | ug/L | 0.96386 | 2011-08-04 09:30 | 20112170170 | Crescent | 06MM136 | Crescent 06MM136: 80 µl | 1.328 | 1-METHYLNAPHTHALENE | . | 0.04 | 1.28 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.96386 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-01-16 | ug/L | 0.9488 | 2011-11-03 07:30 | 20113080004 | Crescent | 06MM136 | Crescent 06MM136: 80 µl | 1.328 | 1-METHYLNAPHTHALENE | . | 0.04 | 1.26 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.9488 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2012-07-27 | ug/L | 1.0316 | 2012-05-17 07:30 | 20121390007 | Crescent | 06MM136 | Crescent 06MM136: 80 µl | 1.328 | 1-METHYLNAPHTHALENE | . | 0.04 | 1.37 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 1.0316 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-06-29 | ug/L | 0.95164 | 2011-06-02 09:00 | 20111540005 | Supelco | LB77171 | Supelco LB77171: 80 µl | 1.3556 | 1-METHYLNAPHTHALENE | . | 0.04 | 1.29 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.95164 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289
1-METHYLNAPHTHALENE | 4433 | 81696GCM99 | 2011-08-07 | ug/L | 0.84895 | 2011-07-14 09:30 | 20111960036 | Supelco | LB77171 | Supelco LB77171: 80 µl | 1.3664 | 1-METHYLNAPHTHALENE | . | 0.04 | 1.16 | MICROGRAMS PER LITER |  |  | >= Reporting Level | 0.84895 | Y | Recoverable Result | Y | [5 empty] | 76 | 0.90088 | 0.12474 | 0.92441 | 0.17324 | 0.12842 | 0.92289