## IBSP Data
Parameter Description | Parmcode | Methcode | Phase | Water Year | Lab Sample ID | Analysis Date | Lab Value | Expected Conc. | % Recov | % Bias | # of QCUs | Report Level | Labcode | Data Qualifier1 | Data Qualifier2 | Data Qualifier3 | Units | Blind Sample ID | Setno
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1190148 | 2015-05-12 | 3.52 | 3.49 | 100.8596 | 0.8596 | 0.12922 |  | 3302 | See result laboratory comment |  |  | ug/L | 25213-75211 | TCMS051115
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1620201 | 2015-07-20 | 4.574 | 4.575 | 99.9781 | -0.021858 | -0.0036349 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS072015
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1740034 | 2015-07-23 | 4.53 | 4.575 | 99.0164 | -0.98361 | -0.16357 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS072215
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2230087 | 2015-09-10 | 4.628 | 4.6087 | 100.4177 | 0.41768 | 0.069634 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25209-75213 | TCMS091015
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2370155 | 2015-09-19 | 4.77 | 4.6087 | 103.4988 | 3.4988 | 0.5833 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25209-75213 | TCMS091815
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1320073 | 2015-06-09 | 0.7861 | 0.75625 | 103.9471 | 3.9471 | 0.199 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | 25211-75209 | TCMS060915
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1470115 | 2015-06-12 | 0.7916 | 0.75625 | 104.6744 | 4.6744 | 0.23567 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | 25211-75209 | TCMS061115
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1250074 | 2015-06-08 | 4.489 | 4.6087 | 97.4017 | -2.5983 | -0.43318 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25209-75213 | TCMS060615
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1390153 | 2015-06-10 | 5.343 | 4.6087 | 115.9317 | 15.9317 | 2.6561 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25209-75213 | TCMS060915
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2160028 | 2015-08-31 | 0.657 | 0.75625 | 86.876 | -13.124 | -0.66167 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | 25211-75209 | TCMS083115
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2320211 | 2015-09-17 | 0.8857 | 0.75625 | 117.1174 | 17.1174 | 0.863 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25211-75209 | TCMS091515
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1880120 | 2015-08-12 | 3.731 | 3.49 | 106.9054 | 6.9054 | 1.0381 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25213-75211 | TCMS081115
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2030107 | 2015-08-24 | 3.471 | 3.49 | 99.4556 | -0.54441 | -0.081839 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 25213-75211 | TCMS082415
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2520016 | 2015-09-29 | 1.794 | 1.965 | 91.2977 | -8.7023 | -0.99532 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75211-25DIW | TCMS092915
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2660034 | 2015-10-07 | 1.807 | 1.965 | 91.9593 | -8.0407 | -0.91965 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75211-25DIW | TCMS100715
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2450074 | 2015-09-25 | 4.631 | 4.575 | 101.224 | 1.224 | 0.20356 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS092415
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2580071 | 2015-10-01 | 4.559 | 4.575 | 99.6503 | -0.34973 | -0.058159 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS093015
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1530193 | 2015-06-12 | 2.114 | 1.965 | 107.5827 | 7.5827 | 0.86726 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75211-25DIW | TCMS061115
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1690074 | 2015-07-22 | 1.961 | 1.965 | 99.7964 | -0.20356 | -0.023282 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75211-25DIW | TCMS072215
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 1950022 | 2015-08-13 | 4.631 | 4.575 | 101.224 | 1.224 | 0.20356 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS081215
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2015 | 2090026 | 2015-08-25 | 4.772 | 4.575 | 104.306 | 4.306 | 0.71608 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS082415
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2790069 | 2015-10-14 | 4.71 | 4.575 | 102.9508 | 2.9508 | 0.49072 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS101315
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2930192 | 2015-11-04 | 4.505 | 4.575 | 98.4699 | -1.5301 | -0.25445 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS110415
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0960078 | 2016-04-22 | 4.708 | 4.575 | 102.9071 | 2.9071 | 0.48345 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCME042116
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1110067 | 2016-04-28 | 4.774 | 4.575 | 104.3497 | 4.3497 | 0.72335 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS042816
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2870054 | 2015-10-27 | 1.365 | 1.4 | 97.5 | -2.5 | -0.175 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T217 | TCMS102315
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3000098 | 2015-11-05 | 1.429 | 1.4 | 102.0714 | 2.0714 | 0.145 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T217 | TCMS110415
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1030064 | 2016-04-27 | 1.433 | 1.4 | 102.3571 | 2.3571 | 0.165 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T217 | TCMS042616
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1170029 | 2016-04-28 | 1.499 | 1.4 | 107.0714 | 7.0714 | 0.495 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T217 | TCMS042816
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1660104 | 2016-07-20 | 2.68 | 2.62 | 102.2901 | 2.2901 | 0.28846 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T211 | TC071116B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1800078 | 2016-07-26 | 2.646 | 2.62 | 100.9924 | 0.99237 | 0.125 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T211 | TC071416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2290123 | 2016-10-08 | 0.54 | 0.341 | 158.3578 | 58.3578 | 0.796 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS16281B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3070095 | 2015-11-17 | 0.879 | 0.341 | 257.7713 | 157.7713 | 2.69 | 0.4 | 3302 | Value verified by rerun, same method | See result laboratory comment |  | µg/L | T219 | TCMS111615
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3210006 | 2015-12-10 | 0.5294 | 0.341 | 155.2493 | 55.2493 | 0.942 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T219 | TCMS120915
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1240108 | 2016-06-02 | 0.5575 | 0.341 | 163.4897 | 63.4897 | 1.0825 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T219 | TC060116
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0400035 | 2016-03-25 | 0.5104 | 0.341 | 149.6774 | 49.6774 | 0.847 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T219 | TCMS032516
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0540101 | 2016-04-09 | 0.6295 | 0.341 | 184.6041 | 84.6041 | 1.4425 | 0.4 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T219 | TCMS040816
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0120026 | 2016-01-21 | 4.863 | 4.575 | 106.2951 | 6.2951 | 1.0469 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS012116
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0260031 | 2016-02-05 | 4.543 | 4.575 | 99.3005 | -0.69945 | -0.11632 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | 75213-25DIW | TCMS020416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0050023 | 2016-01-20 | 1.422 | 1.4 | 101.5714 | 1.5714 | 0.11 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T217 | TCMS012016
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0200028 | 2016-02-04 | 1.461 | 1.4 | 104.3571 | 4.3571 | 0.305 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T217 | TCMS020416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1880119 | 2016-08-20 | 1.483 | 1.4 | 105.9286 | 5.9286 | 0.415 | 0.4 | 3302 |  |  |  | µg/L | T217 | TC081916
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2010272 | 2016-08-09 | 1.467 | 1.4 | 104.7857 | 4.7857 | 0.335 | 0.4 | 3302 |  |  |  | µg/L | T217 | TC080516
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0750071 | 2016-04-21 | 3.2 | 2.8 | 114.2857 | 14.2857 | 2 | 0.4 | 3302 | Value verified by rerun, same method | See result laboratory comment |  | µg/L | T215 | TCMS042016
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0890172 | 2016-04-22 | 2.93 | 2.8 | 104.6429 | 4.6429 | 0.65 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T215 | TCME042116
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2580002 | 2016-11-10 | 2.99 | 2.8 | 106.7857 | 6.7857 | 0.76 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS16291B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2710083 | 2016-11-29 | 3.05 | 2.8 | 108.9286 | 8.9286 | 1 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS16333A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0690046 | 2016-04-20 | 2.934 | 2.62 | 111.9847 | 11.9847 | 1.5096 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T211 | TCMS042016
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 0820024 | 2016-04-22 | 2.734 | 2.62 | 104.3511 | 4.3511 | 0.54808 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T211 | TCME042116
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2520013 | 2016-10-13 | 2.79 | 2.62 | 106.4885 | 6.4885 | 0.68 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS16287A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2640186 | 2016-11-29 | 2.763 | 2.62 | 105.458 | 5.458 | 0.572 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS16333A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3500111 | 2016-01-04 | 3.16 | 2.62 | 120.6107 | 20.6107 | 2.5962 | 0.4 | 3302 | Value verified by rerun, same method | See result laboratory comment |  | µg/L | T211 | TCMS010416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3630067 | 2016-01-05 | 2.621 | 2.62 | 100.0382 | 0.038168 | 0.0048077 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T211 | TCMS010416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3420099 | 2015-12-21 | 3.102 | 2.8 | 110.7857 | 10.7857 | 1.51 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T215 | TCMS122115
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 3560097 | 2016-01-05 | 2.94 | 2.8 | 105 | 5 | 0.7 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T215 | TCMS010416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1590119 | 2016-06-30 | 3.092 | 2.8 | 110.4286 | 10.4286 | 1.46 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T215 | TC062916
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1740273 | 2016-07-26 | 2.824 | 2.8 | 100.8571 | 0.85714 | 0.12 | 0.4 | 3302 | See result laboratory comment |  |  | µg/L | T215 | TC071416
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 1940102 | 2016-08-08 | 4.739 | 4.575 | 103.5847 | 3.5847 | 0.59613 | 0.4 | 3302 |  |  |  | µg/L | 75213-25DIW | TC080516
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2016 | 2080216 | 2016-08-20 | 4.883 | 4.575 | 106.7322 | 6.7322 | 1.1196 | 0.4 | 3302 |  |  |  | µg/L | 75213-25DIW | TC081916
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2780114 | 2016-11-29 | 1.69 | 1.71 | 98.8304 | -1.1696 | -0.08 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS16334B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2940163 | 2016-11-30 | 2.07 | 1.71 | 121.0526 | 21.0526 | 1.44 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS16335A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0940130 | 2017-04-11 | 1.9998 | 1.71 | 116.9474 | 16.9474 | 1.1592 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17101A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1080100 | 2017-05-09 | 1.7327 | 1.71 | 101.3275 | 1.3275 | 0.0908 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17129A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2860091 | 2016-11-29 | 1.6 | 1.4 | 114.2857 | 14.2857 | 0.8 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS16334B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3000017 | 2016-12-05 | 1.67 | 1.4 | 119.2857 | 19.2857 | 1.08 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS16340A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1020077 | 2017-05-10 | 1.4221 | 1.4 | 101.5786 | 1.5786 | 0.0884 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17129A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1150158 | 2017-05-10 | 1.4121 | 1.4 | 100.8643 | 0.86429 | 0.0484 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17129A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3200136 | 2017-01-26 | 0.81 | 0.516 | 156.9767 | 56.9767 | 1.176 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17026A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3370119 | 2016-12-13 | 0.62 | 0.516 | 120.155 | 20.155 | 0.416 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS16348A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1290178 | 2017-06-20 | 0.6412 | 0.516 | 124.2636 | 24.2636 | 0.5008 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17171A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1430168 | 2017-06-22 | 0.6824 | 0.516 | 132.2481 | 32.2481 | 0.6656 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17172C
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3130044 | 2016-12-07 | 1.23 | 0.341 | 360.7038 | 260.7038 | 3.556 | 0.5 | 3302 |  |  |  | µg/L | T219 | ICPMS16342A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3270122 | 2016-12-08 | 0.6 | 0.341 | 175.9531 | 75.9531 | 1.036 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS16343A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1220113 | 2017-05-26 | 0.6056 | 0.341 | 177.5953 | 77.5953 | 1.0584 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS17144B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1370163 | 2017-06-20 | 0.9277 | 0.341 | 272.0528 | 172.0528 | 2.3468 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS17171A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0540066 | 2017-03-27 | 1.0939 | 0.341 | 320.7918 | 220.7918 | 3.0116 | 0.5 | 3302 |  |  |  | µg/L | T219 | ICPMS17086A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0330060 | 2017-02-27 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 |  |  |  | µg/L | T223 | ICPMS17058A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0450117 | 2017-02-27 | 0.5694 | 0.516 | 110.3488 | 10.3488 | 0.2136 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17058A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2200291 | 2017-10-02 | 0.5577 | 0.516 | 108.0814 | 8.0814 | 0.1668 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17275A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2340157 | 2017-10-05 | 0.6645 | 0.516 | 128.7791 | 28.7791 | 0.594 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17277B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0100050 | 2017-02-22 | 1.8591 | 1.71 | 108.7193 | 8.7193 | 0.5964 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17053A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0240086 | 2017-02-22 | 1.8117 | 1.71 | 105.9474 | 5.9474 | 0.4068 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17053A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1920149 | 2017-07-21 | 1.735 | 1.71 | 101.462 | 1.462 | 0.1 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17202B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2060194 | 2017-09-29 | 1.7434 | 1.71 | 101.9532 | 1.9532 | 0.1336 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17272A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0040033 | 2017-02-21 | 1.5875 | 1.4 | 113.3929 | 13.3929 | 0.75 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17052B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0190006 | 2017-02-23 | 1.4423 | 1.4 | 103.0214 | 3.0214 | 0.1692 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17053A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1870005 | 2017-07-20 | 1.4431 | 1.4 | 103.0786 | 3.0786 | 0.1724 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17201A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1990222 | 2017-07-27 | 1.3501 | 1.4 | 96.4357 | -3.5643 | -0.1996 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17207B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0730075 | 2017-03-28 | 2.9471 | 2.8 | 105.2536 | 5.2536 | 0.5884 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS17087A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0870131 | 2017-04-12 | 2.7721 | 2.8 | 99.0036 | -0.99643 | -0.1116 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS17101A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2560128 | 2017-10-18 | 3.1355 | 2.8 | 111.9821 | 11.9821 | 1.342 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS17289C
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2690198 | 2017-12-01 | 2.9352 | 2.8 | 104.8286 | 4.8286 | 0.5408 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS17324A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0660085 | 2017-04-10 | 2.8795 | 2.62 | 109.9046 | 9.9046 | 1.038 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17100A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 0800017 | 2017-03-28 | 2.7328 | 2.62 | 104.3053 | 4.3053 | 0.4512 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17087A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2500006 | 2017-10-16 | 2.8461 | 2.62 | 108.6298 | 8.6298 | 0.9044 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17289B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 2620203 | 2017-12-01 | 2.7886 | 2.62 | 106.4351 | 6.4351 | 0.6744 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17324A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3480033 | 2017-01-31 | 3.05 | 2.62 | 116.4122 | 16.4122 | 1.72 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17031B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3630023 | 2017-02-21 | 2.7299 | 2.62 | 104.1947 | 4.1947 | 0.4396 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17052B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1640178 | 2017-06-30 | 2.5776 | 2.62 | 98.3817 | -1.6183 | -0.1696 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17180A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1780227 | 2017-07-20 | 2.8233 | 2.62 | 107.7595 | 7.7595 | 0.8132 | 0.5 | 3302 |  |  |  | µg/L | T211 | ICPMS17201A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3410125 | 2017-05-08 | 2.951 | 2.8 | 105.3929 | 5.3929 | 0.604 | 0.5 | 3302 | Value verified by rerun, same method |  |  | µg/L | T215 | ICPMS17128A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 3550006 | 2017-01-31 | 4.3 | 2.8 | 153.5714 | 53.5714 | 6 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS17031B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1570092 | 2017-06-27 | 2.8575 | 2.8 | 102.0536 | 2.0536 | 0.23 | 0.5 | 3302 |  |  |  | µg/L | T215 | ICPMS17178A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2017 | 1710241 | 2017-07-19 | 2.8472 | 2.8 | 101.6857 | 1.6857 | 0.1888 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T215 | ICPMS17200B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2760114 | 2017-12-04 | 2.1354 | 1.71 | 124.8772 | 24.8772 | 1.7016 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS17338A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2910001 | 2017-12-13 | 1.8042 | 1.71 | 105.5088 | 5.5088 | 0.3768 | 0.5 | 3302 | Below the reporting level but at or above the detection level | Sample was diluted |  | µg/L | T221 | ICPMS17346B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0930126 | 2018-05-01 | 1.9082 | 1.71 | 111.5906 | 11.5906 | 0.7928 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18121B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1080117 | 2018-05-07 | 1.8223 | 1.71 | 106.5673 | 6.5673 | 0.4492 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18127A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2860038 | 2017-12-13 | 1.7508 | 1.4 | 125.0571 | 25.0571 | 1.4032 | 0.5 | 3302 | Below the reporting level but at or above the detection level | Sample was diluted |  | µg/L | T217 | ICPMS17346B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2970215 | 2017-12-06 | 4.2999 | 1.4 | 307.1357 | 207.1357 | 11.5996 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS17339C
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1020132 | 2018-05-01 | 1.4174 | 1.4 | 101.2429 | 1.2429 | 0.0696 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS18121B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1140191 | 2018-05-08 | 1.7613 | 1.4 | 125.8071 | 25.8071 | 1.4452 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS18127A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 3180171 | 2017-12-13 | 0.7746 | 0.516 | 150.1163 | 50.1163 | 1.0344 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS17346B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 3330004 | 2018-01-10 | 1 | 0.516 | 193.7984 | 93.7984 | 1.936 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T223 | ICPMS18010B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1350128 | 2018-06-06 | 0.6097 | 0.516 | 118.1589 | 18.1589 | 0.3748 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS18157B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1500140 | 2018-06-18 | 0.6428 | 0.516 | 124.5736 | 24.5736 | 0.5072 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS18169A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1280142 | 2018-05-23 | 0.9099 | 0.341 | 266.8328 | 166.8328 | 2.2756 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS18143A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1420223 | 2018-06-08 | 1.46 | 0.341 | 428.1525 | 328.1525 | 4.476 | 0.5 | 3302 |  |  |  | µg/L | T219 | ICPMS18158B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0440095 | 2018-03-06 | 0.6465 | 0.341 | 189.5894 | 89.5894 | 1.222 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS18060A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0580118 | 2018-03-14 | 0.6103 | 0.341 | 178.9736 | 78.9736 | 1.0772 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS18071A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2260149 | 2018-08-28 | 0.7457 | 0.341 | 218.6804 | 118.6804 | 1.6188 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T219 | ICPMS18240A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0370095 | 2018-02-21 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 |  |  |  | µg/L | T223 | ICPMS18051B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0520070 | 2018-03-13 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 |  |  |  | µg/L | T223 | ICPMS18071A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2200026 | 2018-08-21 | 0.6297 | 0.516 | 122.0349 | 22.0349 | 0.4548 | 0.5 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T223 | ICPMS18233A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2330199 | 2018-08-28 | 0.7039 | 0.516 | 136.4147 | 36.4147 | 0.7516 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS18240A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0170071 | 2018-02-01 | 1.7254 | 1.71 | 100.9006 | 0.90058 | 0.0616 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18032A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0300092 | 2018-02-13 | 1.5642 | 1.71 | 91.4737 | -8.5263 | -0.5832 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18039A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1910179 | 2018-07-25 | 1.8771 | 1.71 | 109.7719 | 9.7719 | 0.6684 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T221 | ICPMS18201A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2050299 | 2018-08-03 | 1.8222 | 1.71 | 106.5614 | 6.5614 | 0.4488 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18215B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0100019 | 2018-01-25 | 1.4699 | 1.4 | 104.9929 | 4.9929 | 0.2796 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS18024B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0240044 | 2018-02-13 | 1.3399 | 1.4 | 95.7071 | -4.2929 | -0.2404 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS18039A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1840099 | 2018-07-26 | 1.492 | 1.4 | 106.5714 | 6.5714 | 0.368 | 0.5 | 3302 |  |  |  | µg/L | T217 | ICPMS18201A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1990130 | 2018-08-02 | 1.8132 | 1.4 | 129.5143 | 29.5143 | 1.6528 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T217 | ICPMS18214B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0860122 | 2018-04-16 | 0.5006 | 0.445 | 112.4944 | 12.4944 | 0.2224 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS18106A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0650045 | 2018-03-21 | 2.0202 | 1.17 | 172.6667 | 72.6667 | 3.4008 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18079A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 0790092 | 2018-04-17 | 1.8587 | 1.17 | 158.8632 | 58.8632 | 2.7548 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18106A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2490026 | 2018-09-19 | 1.4777 | 1.17 | 126.2991 | 26.2991 | 1.2308 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T225 | ICPMS18261B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 2610146 | 2018-10-04 | 1.4475 | 1.17 | 123.7179 | 23.7179 | 1.11 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18271B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 3460079 | 2018-01-04 | 1.4547 | 1.17 | 124.3333 | 24.3333 | 1.1388 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18004A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 3610007 | 2018-01-31 | 1.5111 | 1.17 | 129.1538 | 29.1538 | 1.3644 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18031A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1640022 | 2018-07-17 | 1.5808 | 1.17 | 135.1111 | 35.1111 | 1.6432 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18193B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1770196 | 2018-07-20 | 1.4623 | 1.17 | 124.9829 | 24.9829 | 1.1692 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS18198A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 3390056 | 2018-01-03 | 0.6918 | 0.445 | 155.4607 | 55.4607 | 0.9872 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS18003B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1560173 | 2018-07-16 | 0.6942 | 0.445 | 156 | 56 | 0.9968 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS18193B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2018 | 1700199 | 2018-08-03 | 0.6119 | 0.445 | 137.5056 | 37.5056 | 0.6676 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS18215B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2760034 | 2018-10-15 | 1.721 | 1.71 | 100.6433 | 0.64327 | 0.044 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18288B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2890061 | 2018-10-29 | 1.9334 | 1.71 | 113.0643 | 13.0643 | 0.8936 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS18299B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 0920065 | 2019-04-30 | 1.5032 | 1.71 | 87.9064 | -12.0936 | -0.8272 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS19119B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1060103 | 2019-05-06 | 2.1136 | 1.71 | 123.6023 | 23.6023 | 1.6144 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS19122B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2830105 | 2018-10-29 | 7.9198 | 7.51 | 105.4567 | 5.4567 | 1.0245 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS18299B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2960205 | 2018-11-27 | 8.0227 | 7.51 | 106.8269 | 6.8269 | 1.2818 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS18324A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 0990128 | 2019-05-06 | 7.6319 | 7.51 | 101.6232 | 1.6232 | 0.30475 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19122B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1130095 | 2019-05-06 | 7.6586 | 7.51 | 101.9787 | 1.9787 | 0.3715 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS19122B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 3180066 | 2019-02-07 | 1.5 | 0.516 | 290.6977 | 190.6977 | 3.936 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T223 | ICPMS19036B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 3320104 | 2019-02-11 | 0.7135 | 0.516 | 138.2752 | 38.2752 | 0.79 | 0.5 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T223 | ICPMS19042A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1340212 | 2019-06-06 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 |  |  |  | µg/L | T223 | ICPMS19157A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1490062 | 2019-06-24 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 |  |  |  | µg/L | T223 | ICPMS19169A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1410229 | 2019-06-12 | 0.5092 | 0.265 | 192.1509 | 92.1509 | 0.9768 | 0.5 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T231 | ICPMS19158A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 0520007 | 2019-03-29 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T223 | ICPMS19087B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2180150 | 2019-08-29 | 0.879 | 0.516 | 170.3488 | 70.3488 | 1.452 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T223 | ICPMS19239B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2320181 | 2019-09-09 | 0.5 | 0.516 | 96.8992 | -3.1008 | -0.064 | 0.5 | 3302 |  |  |  | µg/L | T223 | ICPMS19252A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1900100 | 2019-08-19 | 2.2734 | 1.71 | 132.9474 | 32.9474 | 2.2536 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T221 | ICPMS19231B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2050046 | 2019-08-19 | 2.2868 | 1.71 | 133.731 | 33.731 | 2.3072 | 0.5 | 3302 |  |  |  | µg/L | T221 | ICPMS19231B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1830106 | 2019-08-15 | 8.2773 | 7.51 | 110.217 | 10.217 | 1.9183 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T229 | ICPMS19219B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2000016 | 2019-08-20 | 7.9505 | 7.51 | 105.8655 | 5.8655 | 1.1013 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19231B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2550075 | 2019-10-01 | 0.6299 | 0.445 | 141.5506 | 41.5506 | 0.7396 | 0.5 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T227 | ICPMS19274A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 0640092 | 2019-04-17 | 1.5964 | 1.17 | 136.4444 | 36.4444 | 1.7056 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T225 | ICPMS19105A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 0780096 | 2019-04-26 | 2.2527 | 1.17 | 192.5385 | 92.5385 | 4.3308 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T225 | ICPMS19114A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2470180 | 2019-10-02 | 1.882 | 1.17 | 160.8547 | 60.8547 | 2.848 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS19274A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 2600328 | 2019-10-09 | 2.1193 | 1.17 | 181.1368 | 81.1368 | 3.7972 | 0.5 | 3302 | Analyte detected in laboratory blank | Sample was diluted | See result laboratory comment | µg/L | T225 | ICPMS19282B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 3450114 | 2019-02-21 | 1.9203 | 1.17 | 164.1282 | 64.1282 | 3.0012 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS19045B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1620136 | 2019-07-03 | 1.6582 | 1.17 | 141.7265 | 41.7265 | 1.9528 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS19183B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1760173 | 2019-07-23 | 1.7417 | 1.17 | 148.8632 | 48.8632 | 2.2868 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS19203A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 3400024 | 2019-03-06 | 0.7055 | 0.445 | 158.5393 | 58.5393 | 1.042 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS19065A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 3530016 | 2019-02-04 | 0.7701 | 0.445 | 173.0562 | 73.0562 | 1.3004 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS19035A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1550162 | 2019-06-26 | 0.6319 | 0.445 | 142 | 42 | 0.7476 | 0.5 | 3302 | Below the reporting level but at or above the detection level | See result laboratory comment |  | µg/L | T227 | ICPMS19175B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2019 | 1690137 | 2019-07-02 | 0.6943 | 0.445 | 156.0225 | 56.0225 | 0.9972 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T227 | ICPMS19183B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2810114 | 2019-12-19 | 1.5626 | 1.86 | 84.0108 | -15.9892 | -1.1896 | 0.5 | 3302 | Analyte detected in laboratory blank | See result laboratory comment |  | µg/L | T233 | ICPMS19352B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2950100 | 2019-12-17 | 1.7946 | 1.86 | 96.4839 | -3.5161 | -0.2616 | 0.5 | 3302 | Analyte detected in laboratory blank | See result laboratory comment |  | µg/L | T233 | ICPMS19347B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0980043 | 2020-04-23 | 1.7444 | 1.86 | 93.7849 | -6.2151 | -0.4624 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20114A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1120049 | 2020-05-04 | 1.7628 | 1.86 | 94.7742 | -5.2258 | -0.3888 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20125B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2890221 | 2019-12-12 | 7.0708 | 7.51 | 94.1518 | -5.8482 | -1.098 | 0.5 | 3302 | Analyte detected in laboratory blank | See result laboratory comment |  | µg/L | T229 | ICPMS19344B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 3030027 | 2019-12-18 | 7.5713 | 7.51 | 100.8162 | 0.81625 | 0.15325 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19350A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1060059 | 2020-04-23 | 6.9117 | 7.51 | 92.0333 | -7.9667 | -1.4958 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T229 | ICPMS20114A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 3170095 | 2019-12-30 | 1.8864 | 1.88 | 100.3404 | 0.34043 | 0.0256 | 0.5 | 3302 | Analyte detected in laboratory blank | See result laboratory comment |  | µg/L | T235 | ICPMS19364A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 3310060 | 2019-12-30 | 1.7326 | 1.88 | 92.1596 | -7.8404 | -0.5896 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T235 | ICPMS19364A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1330126 | 2020-06-12 | 1.789 | 1.88 | 95.1596 | -4.8404 | -0.364 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS20163B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0350135 | 2020-02-12 | 1.8751 | 1.88 | 99.7394 | -0.26064 | -0.0196 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T235 | ICPMS20043A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0500059 | 2020-03-04 | 1.6212 | 1.88 | 86.234 | -13.766 | -1.0352 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T235 | ICPMS20064A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2170035 | 2020-09-17 | 1.7953 | 1.88 | 95.4947 | -4.5053 | -0.3388 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS20260A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2320037 | 2020-09-24 | 1.6353 | 1.88 | 86.984 | -13.016 | -0.9788 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS20276A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0150017 | 2020-02-03 | 1.8899 | 1.86 | 101.6075 | 1.6075 | 0.1196 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20031A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0280070 | 2020-02-03 | 1.7531 | 1.86 | 94.2527 | -5.7473 | -0.4276 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20031A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1960126 | 2020-09-03 | 1.6431 | 1.86 | 88.3387 | -11.6613 | -0.8676 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20247B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2120199 | 2020-09-17 | 1.9001 | 1.86 | 102.1559 | 2.1559 | 0.1604 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20260A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0070044 | 2020-01-21 | 7.1927 | 7.51 | 95.775 | -4.225 | -0.79325 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T229 | ICPMS20021A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0220109 | 2020-02-03 | 7.3306 | 7.51 | 97.6112 | -2.3888 | -0.4485 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T229 | ICPMS20031A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1890182 | 2020-09-02 | 7.4248 | 7.51 | 98.8655 | -1.1345 | -0.213 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS20245B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2040025 | 2020-09-03 | 7.0311 | 7.51 | 93.6232 | -6.3768 | -1.1972 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS20247B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0630051 | 2020-03-18 | 1.3626 | 1.17 | 116.4615 | 16.4615 | 0.7704 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T225 | ICPMS20076A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 0770114 | 2020-06-03 | 1.0758 | 1.17 | 91.9487 | -8.0513 | -0.3768 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS20153B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2530137 | 2020-10-06 | 1.2455 | 1.17 | 106.453 | 6.453 | 0.302 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS20279B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 2670298 | 2020-10-09 | 1.3003 | 1.17 | 111.1368 | 11.1368 | 0.5212 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS20281B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 3440294 | 2020-01-30 | 1.1989 | 1.17 | 102.4701 | 2.4701 | 0.1156 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS20030A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 3600045 | 2020-01-21 | 1.1795 | 1.17 | 100.812 | 0.81197 | 0.038 | 0.5 | 3302 | See result laboratory comment |  |  | µg/L | T225 | ICPMS20021A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1610191 | 2020-07-01 | 1.1982 | 1.17 | 102.4103 | 2.4103 | 0.1128 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS20182B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2020 | 1760101 | 2020-08-28 | 1.1328 | 1.17 | 96.8205 | -3.1795 | -0.1488 | 0.5 | 3302 |  |  |  | µg/L | T225 | ICPMS20240A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2810032 | 2020-10-22 | 1.8655 | 1.86 | 100.2957 | 0.2957 | 0.022 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS20296B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2950028 | 2021-01-13 | 1.4261 | 1.86 | 76.672 | -23.328 | -1.7356 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21013B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0970048 | 2021-05-21 | 1.6456 | 1.86 | 88.4731 | -11.5269 | -0.8576 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21140A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1110004 | 2021-06-02 | 1.7219 | 1.86 | 92.5753 | -7.4247 | -0.5524 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21153A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2880069 | 2020-12-30 | 7.6209 | 7.51 | 101.4767 | 1.4767 | 0.27725 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS20350A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3010155 | 2021-01-14 | 7.5326 | 7.51 | 100.3009 | 0.30093 | 0.0565 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21013B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1030144 | 2021-05-05 | 7.3384 | 7.51 | 97.715 | -2.285 | -0.429 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21125A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1180038 | 2021-06-16 | 7.2131 | 7.51 | 96.0466 | -3.9534 | -0.74225 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21167B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3150100 | 2021-01-15 | 1.5171 | 1.88 | 80.6968 | -19.3032 | -1.4516 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21015A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3300047 | 2021-02-02 | 1.8329 | 1.88 | 97.4947 | -2.5053 | -0.1884 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21033A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1320162 | 2021-06-29 | 1.7098 | 1.88 | 90.9468 | -9.0532 | -0.6808 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21180B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1460003 | 2021-07-10 | 1.5633 | 1.88 | 83.1543 | -16.8457 | -1.2668 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21188B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0480038 | 2021-05-18 | 0.8717 | 0.265 | 328.9434 | 228.9434 | 2.4268 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T231 | ICPMS21138A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2220178 | 2021-09-23 | 0.5014 | 0.265 | 189.2075 | 89.2075 | 0.9456 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T231 | ICPMS21258B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0400076 | 2021-05-04 | 1.7606 | 1.88 | 93.6489 | -6.3511 | -0.4776 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21124A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0540055 | 2021-05-11 | 1.8936 | 1.88 | 100.7234 | 0.7234 | 0.0544 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21131A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2150046 | 2021-08-30 | 1.8773 | 1.88 | 99.8564 | -0.14362 | -0.0108 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21242B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2290121 | 2021-09-24 | 1.8365 | 1.88 | 97.6862 | -2.3138 | -0.174 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS21267A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0120036 | 2021-03-04 | 1.7775 | 1.86 | 95.5645 | -4.4355 | -0.33 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21055A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0270078 | 2021-03-26 | 1.5954 | 1.86 | 85.7742 | -14.2258 | -1.0584 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21077A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1940229 | 2021-08-16 | 1.8657 | 1.86 | 100.3065 | 0.30645 | 0.0228 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21228B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2080172 | 2021-08-19 | 1.7611 | 1.86 | 94.6828 | -5.3172 | -0.3956 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21230A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0050024 | 2021-03-04 | 7.1768 | 7.51 | 95.5632 | -4.4368 | -0.833 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21055A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0210015 | 2021-03-26 | 6.8976 | 7.51 | 91.8455 | -8.1545 | -1.531 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21077A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1880114 | 2021-08-11 | 7.1017 | 7.51 | 94.5632 | -5.4368 | -1.0207 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21222B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2010231 | 2021-08-19 | 7.0291 | 7.51 | 93.5965 | -6.4035 | -1.2023 | 0.5 | 3302 |  |  |  | µg/L | T229 | ICPMS21230A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0680063 | 2021-05-14 | 3.3605 | 3.06 | 109.8203 | 9.8203 | 1.202 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T239 | ICPMS21132B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0830027 | 2021-05-15 | 3.5213 | 3.06 | 115.0752 | 15.0752 | 1.8452 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T239 | ICPMS21134A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2580265 | 2021-10-20 | 3.0821 | 3.06 | 100.7222 | 0.72222 | 0.0884 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS21292A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2790034 | 2021-11-23 | 3.0369 | 3.06 | 99.2451 | -0.7549 | -0.0924 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS21308B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0620100 | 2021-05-13 | 2.8139 | 2.75 | 102.3236 | 2.3236 | 0.2556 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21132B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 0760035 | 2021-05-14 | 2.7066 | 2.75 | 98.4218 | -1.5782 | -0.1736 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21134A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2510433 | 2021-10-14 | 2.8133 | 2.75 | 102.3018 | 2.3018 | 0.2532 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21287B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 2640271 | 2021-11-02 | 2.6338 | 2.75 | 95.7745 | -4.2255 | -0.4648 | 0.5 | 3302 | Below the reporting level but at or above the detection level | Sample was diluted |  | µg/L | T237 | ICPMS21295B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3500073 | 2021-01-29 | 2.651 | 2.75 | 96.4 | -3.6 | -0.396 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21028B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3640015 | 2021-03-04 | 2.7261 | 2.75 | 99.1309 | -0.86909 | -0.0956 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21055A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1600035 | 2021-07-24 | 2.6368 | 2.75 | 95.8836 | -4.1164 | -0.4528 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21200B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1730218 | 2021-08-07 | 3.0334 | 2.75 | 110.3055 | 10.3055 | 1.1336 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS21218A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3440081 | 2021-01-29 | 2.4728 | 3.06 | 80.8105 | -19.1895 | -2.3488 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS21028B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 3580059 | 2021-03-04 | 2.9946 | 3.06 | 97.8627 | -2.1373 | -0.2616 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS21055A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1540139 | 2021-07-16 | 3.0096 | 3.06 | 98.3529 | -1.6471 | -0.2016 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS21195A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2021 | 1660119 | 2021-07-29 | 2.4997 | 3.06 | 81.6895 | -18.3105 | -2.2412 | 0.5 | 3302 | Sample was diluted |  |  | µg/L | T239 | ICPMS21209B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2790037 | 2021-12-14 | 1.7562 | 1.86 | 94.4194 | -5.5806 | -0.4152 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS21340B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0960123 | 2022-05-16 | 2 | 1.86 | 107.5269 | 7.5269 | 0.56 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS22136A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1090127 | 2022-05-19 | 1.9 | 1.86 | 102.1505 | 2.1505 | 0.16 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS22139A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3130184 | 2022-01-05 | 2.0187 | 1.88 | 107.3777 | 7.3777 | 0.5548 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22005A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3270288 | 2022-01-20 | 1.9 | 1.88 | 101.0638 | 1.0638 | 0.08 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22011B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1320258 | 2022-07-20 | 2.1 | 1.88 | 111.7021 | 11.7021 | 0.88 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22157B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1460098 | 2022-06-27 | 1.9 | 1.88 | 101.0638 | 1.0638 | 0.08 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22178A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3060209 | 2022-01-04 | 0.9599 | 0.945 | 101.5767 | 1.5767 | 0.0596 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS21356B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3200088 | 2022-01-06 | 1 | 0.945 | 105.8201 | 5.8201 | 0.22 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22005A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1230094 | 2022-05-26 | 1 | 0.945 | 105.8201 | 5.8201 | 0.22 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22144A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1380070 | 2022-06-24 | 0.9 | 0.945 | 95.2381 | -4.7619 | -0.18 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22174A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0390071 | 2022-03-18 | 0.8 | 0.945 | 84.6561 | -15.3439 | -0.58 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22076B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0540113 | 2022-04-20 | 0.9 | 0.945 | 95.2381 | -4.7619 | -0.18 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22088B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2210114 | 2022-08-19 | 0.8 | 0.945 | 84.6561 | -15.3439 | -0.58 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22229B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2360390 | 2022-09-21 | 0.8 | 0.945 | 84.6561 | -15.3439 | -0.58 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS22263B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0380065 | 2022-03-17 | 1.8 | 1.88 | 95.7447 | -4.2553 | -0.32 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22076B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0480131 | 2022-03-24 | 1.8 | 1.88 | 95.7447 | -4.2553 | -0.32 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22080A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2140174 | 2022-08-09 | 1.7 | 1.88 | 90.4255 | -9.5745 | -0.72 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22220A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2280351 | 2022-08-26 | 1.9 | 1.88 | 101.0638 | 1.0638 | 0.08 | 0.5 | 3302 |  |  |  | µg/L | T235 | ICPMS22238A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0120120 | 2022-02-17 | 2.1 | 1.86 | 112.9032 | 12.9032 | 0.96 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS22047B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0250044 | 2022-03-02 | 2 | 1.86 | 107.5269 | 7.5269 | 0.56 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS22060A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1930200 | 2022-08-04 | 1.9 | 1.86 | 102.1505 | 2.1505 | 0.16 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS22215B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2080416 | 2022-08-08 | 2 | 1.86 | 107.5269 | 7.5269 | 0.56 | 0.5 | 3302 |  |  |  | µg/L | T233 | ICPMS22220A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0670123 | 2022-04-20 | 3 | 3.06 | 98.0392 | -1.9608 | -0.24 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22088B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0820073 | 2022-05-24 | 3.1 | 3.06 | 101.3072 | 1.3072 | 0.16 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22143B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2570117 | 2022-09-28 | 3 | 3.06 | 98.0392 | -1.9608 | -0.24 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22269A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0760093 | 2022-05-24 | 3.1 | 2.75 | 112.7273 | 12.7273 | 1.4 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22143B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 0880096 | 2022-07-29 | 2.7 | 2.75 | 98.1818 | -1.8182 | -0.2 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22172A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2500494 | 2022-10-05 | 2.6 | 2.75 | 94.5455 | -5.4545 | -0.6 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22277A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 2630184 | 2022-10-05 | 3.6 | 2.75 | 130.9091 | 30.9091 | 3.4 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22277A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3480118 | 2022-01-21 | 2.7 | 2.75 | 98.1818 | -1.8182 | -0.2 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22011B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3640134 | 2022-01-29 | 2.7 | 2.75 | 98.1818 | -1.8182 | -0.2 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22027D
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1660162 | 2022-07-26 | 2.6 | 2.75 | 94.5455 | -5.4545 | -0.6 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22203B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1790078 | 2022-07-28 | 2.7 | 2.75 | 98.1818 | -1.8182 | -0.2 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS22208B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3410064 | 2022-01-15 | 2.9 | 3.06 | 94.7712 | -5.2288 | -0.64 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22014B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 3570003 | 2022-01-28 | 2.7 | 3.06 | 88.2353 | -11.7647 | -1.44 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22027C
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1580091 | 2022-07-22 | 3.1 | 3.06 | 101.3072 | 1.3072 | 0.16 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22202A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2022 | 1730126 | 2022-07-28 | 3.2 | 3.06 | 104.5752 | 4.5752 | 0.56 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS22208B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2770114 | 2022-10-26 | 2.2 | 2.36 | 93.2203 | -6.7797 | -0.64 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS22298B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2920222 | 2022-10-28 | 2.4 | 2.36 | 101.6949 | 1.6949 | 0.16 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS22300A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0940130 | 2023-04-28 | 2.5408 | 2.36 | 107.661 | 7.661 | 0.7232 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS23117A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1100089 | 2023-05-01 | 2.5107 | 2.36 | 106.3856 | 6.3856 | 0.6028 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS23121B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3130109 | 2022-11-16 | 7.599 | 8.2 | 92.6707 | -7.3293 | -1.047 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS22320A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3330054 | 2023-02-02 | 7.7472 | 8.2 | 94.478 | -5.522 | -0.78885 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23032A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1300167 | 2023-05-24 | 7.2214 | 8.2 | 88.0659 | -11.9341 | -1.7049 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23142B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1440060 | 2023-06-06 | 7.5072 | 8.2 | 91.5512 | -8.4488 | -1.207 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23156B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3060049 | 2022-11-08 | 1.017 | 0.945 | 107.619 | 7.619 | 0.288 | 0.5 | 3302 |  |  |  | µg/L | T243 | ICPMS22312B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3200145 | 2023-01-28 | 0.8082 | 0.945 | 85.5238 | -14.4762 | -0.5472 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS23027A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1240210 | 2023-05-20 | 0.9178 | 0.945 | 97.1217 | -2.8783 | -0.1088 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS23139A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1360048 | 2023-05-25 | 0.837 | 0.945 | 88.5714 | -11.4286 | -0.432 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS23144C
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0470007 | 2023-03-16 | 0.8811 | 0.945 | 93.2381 | -6.7619 | -0.2556 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS23074A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0620097 | 2023-03-25 | 0.8737 | 0.945 | 92.455 | -7.545 | -0.2852 | 0.5 | 3302 | Below the reporting level but at or above the detection level |  |  | µg/L | T243 | ICPMS23083B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2280173 | 2023-09-08 | 1.1284 | 0.945 | 119.4074 | 19.4074 | 0.7336 | 0.5 | 3302 |  |  |  | µg/L | T243 | ICPMS23250A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0380151 | 2023-02-18 | 7.4149 | 8.2 | 90.4256 | -9.5744 | -1.3678 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23048A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0590092 | 2023-03-16 | 7.3851 | 8.2 | 90.0622 | -9.9378 | -1.4197 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23074A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2200127 | 2023-09-01 | 7.372 | 8.2 | 89.9024 | -10.0976 | -1.4425 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23241B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2350096 | 2023-09-27 | 7.4543 | 8.2 | 90.9061 | -9.0939 | -1.2991 | 0.5 | 3302 |  |  |  | µg/L | T247 | ICPMS23269A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0190114 | 2023-02-16 | 2.5022 | 2.36 | 106.0254 | 6.0254 | 0.5688 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS23047A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0310087 | 2023-02-17 | 2.0959 | 2.36 | 88.8093 | -11.1907 | -1.0564 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS23048A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1940234 | 2023-07-28 | 2.3697 | 2.36 | 100.411 | 0.41102 | 0.0388 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS23208B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2060142 | 2023-08-08 | 2.5811 | 2.36 | 109.3686 | 9.3686 | 0.8844 | 0.5 | 3302 |  |  |  | µg/L | T245 | ICPMS23219A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0740033 | 2023-03-22 | 2.9509 | 3.06 | 96.4346 | -3.5654 | -0.4364 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS23080A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0880093 | 2023-05-04 | 2.9357 | 3.06 | 95.9379 | -4.0621 | -0.4972 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS23123B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0670102 | 2023-03-25 | 2.7043 | 2.75 | 98.3382 | -1.6618 | -0.1828 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23083B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0800075 | 2023-03-30 | 2.8705 | 2.75 | 104.3818 | 4.3818 | 0.482 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23089B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 2490138 | 2023-09-18 | 2.6675 | 2.75 | 97 | -3 | -0.33 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23257A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3490121 | 2023-02-09 | 2.7206 | 2.75 | 98.9309 | -1.0691 | -0.1176 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23040A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 0040114 | 2023-02-14 | 3.047 | 2.75 | 110.8 | 10.8 | 1.188 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23045A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1650066 | 2023-06-26 | 2.773 | 2.75 | 100.8364 | 0.83636 | 0.092 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23177B
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1840086 | 2023-07-18 | 2.7843 | 2.75 | 101.2473 | 1.2473 | 0.1372 | 0.5 | 3302 |  |  |  | µg/L | T237 | ICPMS23198A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3400090 | 2023-02-02 | 3.1043 | 3.06 | 101.4477 | 1.4477 | 0.1772 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS23032A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 3610135 | 2023-01-09 | 2.7917 | 3.06 | 91.232 | -8.768 | -1.0732 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS23009A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1570041 | 2023-06-15 | 3.0428 | 3.06 | 99.4379 | -0.56209 | -0.0688 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS23164A
Chromium, UNF, cICP-MS | 01034 | PLM77 | UNF | 2023 | 1730211 | 2023-06-28 | 2.9793 | 3.06 | 97.3627 | -2.6373 | -0.3228 | 0.5 | 3302 |  |  |  | µg/L | T239 | ICPMS23179A