## IBSP Data
Parameter Description | Parmcode | Methcode | Phase | Water Year | Lab Sample ID | Analysis Date | Remark Code | Lab Value | Expected Conc. | % Recovery | % Bias | 1 QCU | # of QCUs | Report Level | Labcode | Value Qualifier1 | Value Qualifier2 | Value Qualifier3 | Units | Blind Sample ID | Setno
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2018 | 2610146 | 2018-10-04 |  | 204.5224 | 183 | 111.7609 | 11.7609 | 19.9 | 1.0815 |  | 3415 |  |  |  | ug/L | T225 | ICPMS18271B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2760034 | 2018-10-15 |  | 17.4469 | 19.2 | 90.8693 | -9.1307 | 2.5 | -0.70124 | 5 | 3415 |  |  |  | µg/L | T221 | ICPMS18288B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2890061 | 2018-10-29 |  | 19.3725 | 19.2 | 100.8984 | 0.89844 | 2.5 | 0.069 | 5 | 3415 |  |  |  | µg/L | T221 | ICPMS18299B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2830105 | 2018-10-29 |  | 108.1835 | 117 | 92.4645 | -7.5355 | 11.3 | -0.78022 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS18299B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2960205 | 2018-11-27 |  | 115.2728 | 117 | 98.5238 | -1.4762 | 11.3 | -0.15285 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS18324A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 3100174 | 2019-02-05 |  | 23.5179 | 25.7 | 91.5093 | -8.4907 | 2.5 | -0.87284 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS19035A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 3320104 | 2019-02-11 |  | 13.9508 | 13 | 107.3138 | 7.3138 | 2.5 | 0.38032 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T223 | ICPMS19042A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 3270026 | 2019-02-12 |  | 24.1674 | 25.7 | 94.0366 | -5.9634 | 2.5 | -0.61304 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS19042A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 3450114 | 2019-02-21 |  | 204.2472 | 183 | 111.6105 | 11.6105 | 19.9 | 1.0677 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS19045B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 0520007 | 2019-03-29 |  | 13.4261 | 13 | 103.2777 | 3.2777 | 2.5 | 0.17044 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T223 | ICPMS19087B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 0570186 | 2019-04-08 |  | 23.6616 | 25.7 | 92.0685 | -7.9315 | 2.5 | -0.81536 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T231 | ICPMS19098A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 0640092 | 2019-04-17 |  | 216.483 | 183 | 118.2967 | 18.2967 | 19.9 | 1.6826 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T225 | ICPMS19105A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 0780096 | 2019-04-26 |  | 233.8256 | 183 | 127.7736 | 27.7736 | 19.9 | 2.5541 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T225 | ICPMS19114A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 0920065 | 2019-04-30 |  | 19.7063 | 19.2 | 102.637 | 2.637 | 2.5 | 0.20252 | 5 | 3415 |  |  |  | µg/L | T221 | ICPMS19119B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1060103 | 2019-05-06 |  | 17.8646 | 19.2 | 93.0448 | -6.9552 | 2.5 | -0.53416 | 5 | 3415 |  |  |  | µg/L | T221 | ICPMS19122B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 0990128 | 2019-06-04 |  | 112.7436 | 117 | 96.3621 | -3.6379 | 11.3 | -0.37667 | 5 | 3415 | Sample was diluted |  |  | µg/L | T229 | ICPMS19149A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1130095 | 2019-06-04 |  | 113.8768 | 117 | 97.3306 | -2.6694 | 11.3 | -0.27639 | 5 | 3415 | Sample was diluted |  |  | µg/L | T229 | ICPMS19149A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1280117 | 2019-06-05 |  | 26.557 | 25.7 | 103.3346 | 3.3346 | 2.5 | 0.3428 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T231 | ICPMS19154A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1340212 | 2019-06-06 |  | 12.6927 | 13 | 97.6362 | -2.3638 | 2.5 | -0.12292 | 5 | 3415 |  |  |  | µg/L | T223 | ICPMS19157A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1410229 | 2019-06-12 |  | 27.164 | 25.7 | 105.6965 | 5.6965 | 2.5 | 0.5856 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T231 | ICPMS19158A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1490062 | 2019-06-24 |  | 13.5123 | 13 | 103.9408 | 3.9408 | 2.5 | 0.20492 | 5 | 3415 |  |  |  | µg/L | T223 | ICPMS19169A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1760173 | 2019-07-03 |  | 219.6227 | 183 | 120.0124 | 20.0124 | 19.9 | 1.8403 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS19184B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1620136 | 2019-07-23 |  | 209.7849 | 183 | 114.6366 | 14.6366 | 19.9 | 1.346 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS19203A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1830106 | 2019-08-02 |  | 107.9793 | 117 | 92.29 | -7.71 | 11.3 | -0.79829 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19214A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 1900100 | 2019-08-19 |  | 18.6571 | 19.2 | 97.1724 | -2.8276 | 2.5 | -0.21716 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T221 | ICPMS19231B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2050046 | 2019-08-19 |  | 18.6667 | 19.2 | 97.2224 | -2.7776 | 2.5 | -0.21332 | 5 | 3415 |  |  |  | µg/L | T221 | ICPMS19231B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2000016 | 2019-08-20 |  | 115.3973 | 117 | 98.6302 | -1.3698 | 11.3 | -0.14183 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19231B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2180150 | 2019-08-29 |  | 12.8713 | 13 | 99.01 | -0.99 | 2.5 | -0.05148 | 5 | 3415 |  |  |  | µg/L | T223 | ICPMS19239B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2250166 | 2019-09-04 |  | 25.4642 | 25.7 | 99.0825 | -0.91751 | 2.5 | -0.09432 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS19247A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2320181 | 2019-09-09 |  | 12.8345 | 13 | 98.7269 | -1.2731 | 2.5 | -0.0662 | 5 | 3415 |  |  |  | µg/L | T223 | ICPMS19252A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2390203 | 2019-09-23 |  | 27.031 | 25.7 | 105.179 | 5.179 | 2.5 | 0.5324 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS19266A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2470180 | 2019-10-09 |  | 213.4638 | 183 | 116.6469 | 16.6469 | 19.9 | 1.5308 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS19282B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2019 | 2600328 | 2019-10-09 |  | 207.9956 | 183 | 113.6588 | 13.6588 | 19.9 | 1.2561 | 5 | 3415 | Sample was diluted |  |  | µg/L | T225 | ICPMS19282B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2810114 | 2019-12-10 |  | 19.1946 | 22 | 87.2482 | -12.7518 | 2.5 | -1.1222 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS19338A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2890221 | 2019-12-12 |  | 120.1536 | 117 | 102.6954 | 2.6954 | 11.3 | 0.27908 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19344B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2950100 | 2019-12-16 |  | 20.8243 | 22 | 94.6559 | -5.3441 | 2.5 | -0.47028 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T233 | ICPMS19347B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3030027 | 2019-12-18 |  | 135.4348 | 117 | 115.7562 | 15.7562 | 11.3 | 1.6314 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS19350A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3090147 | 2019-12-30 |  | 27.9474 | 25.7 | 108.7447 | 8.7447 | 2.5 | 0.89896 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T231 | ICPMS19364A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3230099 | 2019-12-30 |  | 29.2 | 25.7 | 113.6187 | 13.6187 | 2.5 | 1.4 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS19364A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3170095 | 2019-12-30 |  | 21.1236 | 20.7 | 102.0464 | 2.0464 | 2.5 | 0.16944 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T235 | ICPMS19364A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3310060 | 2019-12-30 |  | 22.0438 | 20.7 | 106.4918 | 6.4918 | 2.5 | 0.53752 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T235 | ICPMS19364A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3600045 | 2020-01-21 |  | 226.1959 | 183 | 123.6043 | 23.6043 | 19.9 | 2.1706 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T225 | ICPMS20021A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0070044 | 2020-01-21 |  | 119.0907 | 117 | 101.7869 | 1.7869 | 11.3 | 0.18502 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS20021A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 3440294 | 2020-01-30 |  | 227.1084 | 183 | 124.103 | 24.103 | 19.9 | 2.2165 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS20030A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0220109 | 2020-02-03 |  | 116.1911 | 117 | 99.3086 | -0.69137 | 11.3 | -0.071584 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS20031A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0150017 | 2020-02-03 |  | 20.5754 | 22 | 93.5245 | -6.4755 | 2.5 | -0.56984 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20031A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0280070 | 2020-02-03 |  | 21.1639 | 22 | 96.1995 | -3.8005 | 2.5 | -0.33444 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20031A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0350135 | 2020-02-12 |  | 20.6975 | 20.7 | 99.9879 | -0.012077 | 2.5 | -0.001 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T235 | ICPMS20043A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0420076 | 2020-03-03 |  | 26.9449 | 25.7 | 104.844 | 4.844 | 2.5 | 0.49796 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS20063B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0560053 | 2020-03-04 |  | 25.2711 | 25.7 | 98.3311 | -1.6689 | 2.5 | -0.17156 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T231 | ICPMS20064A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0500059 | 2020-03-04 |  | 19.4674 | 20.7 | 94.0454 | -5.9546 | 2.5 | -0.49304 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T235 | ICPMS20064A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1060059 | 2020-04-23 |  | 108.9749 | 117 | 93.1409 | -6.8591 | 11.3 | -0.71019 | 5 | 3415 | See result laboratory comment |  |  | µg/L | T229 | ICPMS20114A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0980043 | 2020-04-23 |  | 20.0246 | 22 | 91.0209 | -8.9791 | 2.5 | -0.79016 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20114A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1120049 | 2020-05-04 |  | 20.7271 | 22 | 94.2141 | -5.7859 | 2.5 | -0.50916 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20125B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0630051 | 2020-06-03 |  | 208.0213 | 183 | 113.6728 | 13.6728 | 19.9 | 1.2574 | 5 | 3415 | Sample was diluted |  |  | µg/L | T225 | ICPMS20153B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 0770114 | 2020-06-03 |  | 215.1515 | 183 | 117.5691 | 17.5691 | 19.9 | 1.6157 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS20153B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1610191 | 2020-07-01 |  | 206.8356 | 183 | 113.0249 | 13.0249 | 19.9 | 1.1978 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS20182B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1330126 | 2020-08-27 |  | 20.0139 | 20.7 | 96.6855 | -3.3145 | 2.5 | -0.27444 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS20238B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1890182 | 2020-09-02 |  | 117.9838 | 117 | 100.8409 | 0.84085 | 11.3 | 0.087062 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS20245B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2040025 | 2020-09-03 |  | 104.5644 | 117 | 89.3713 | -10.6287 | 11.3 | -1.1005 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS20247B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1960126 | 2020-09-03 |  | 22.5527 | 22 | 102.5123 | 2.5123 | 2.5 | 0.22108 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20247B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2120199 | 2020-09-17 |  | 22.1396 | 22 | 100.6345 | 0.63455 | 2.5 | 0.05584 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20260A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2170035 | 2020-09-17 |  | 21.6977 | 20.7 | 104.8198 | 4.8198 | 2.5 | 0.39908 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS20260A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2240092 | 2020-09-24 |  | 27.5346 | 25.7 | 107.1385 | 7.1385 | 2.5 | 0.73384 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS20276A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2320037 | 2020-09-24 |  | 24.191 | 20.7 | 116.8647 | 16.8647 | 2.5 | 1.3964 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS20276A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2390027 | 2020-09-29 |  | 24.7789 | 25.7 | 96.416 | -3.584 | 2.5 | -0.36844 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS20273A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2670298 | 2020-10-09 |  | 221.0955 | 183 | 120.8172 | 20.8172 | 19.9 | 1.9143 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS20281B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 1760101 | 2020-10-14 |  | 210.0974 | 183 | 114.8073 | 14.8073 | 19.9 | 1.3617 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS20283A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2810032 | 2020-10-22 |  | 19.8782 | 22 | 90.3555 | -9.6445 | 2.5 | -0.84872 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS20296B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2880069 | 2020-12-30 |  | 118.2634 | 117 | 101.0798 | 1.0798 | 11.3 | 0.11181 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS20350A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2020 | 2530137 | 2021-01-12 |  | 182.8664 | 183 | 99.927 | -0.073005 | 19.9 | -0.0067136 | 5 | 3415 |  |  |  | µg/L | T225 | ICPMS20353A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2950028 | 2021-01-13 |  | 19.3548 | 22 | 87.9764 | -12.0236 | 2.5 | -1.0581 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21013B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3010155 | 2021-01-14 |  | 119.7677 | 117 | 102.3656 | 2.3656 | 11.3 | 0.24493 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS21013B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3080107 | 2021-01-14 |  | 26.2034 | 25.7 | 101.9588 | 1.9588 | 2.5 | 0.20136 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS21013B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3150100 | 2021-01-15 |  | 20.8281 | 20.7 | 100.6188 | 0.61884 | 2.5 | 0.05124 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21015A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3230039 | 2021-01-27 |  | 23.6208 | 25.7 | 91.9097 | -8.0903 | 2.5 | -0.83168 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS21026B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3500073 | 2021-01-29 |  | 10.8743 | 10.5 | 103.5648 | 3.5648 | 2.5 | 0.14972 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS21028B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3440081 | 2021-01-29 |  | 218.2068 | 231 | 94.4618 | -5.5382 | 16.8 | -0.7615 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS21028B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3300047 | 2021-02-02 |  | 21.9996 | 20.7 | 106.2783 | 6.2783 | 2.5 | 0.51984 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21033A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0050024 | 2021-03-04 |  | 106.6155 | 117 | 91.1244 | -8.8756 | 11.3 | -0.91898 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS21055A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0120036 | 2021-03-04 |  | 18.8707 | 22 | 85.7759 | -14.2241 | 2.5 | -1.2517 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21055A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3640015 | 2021-03-04 |  | 9.2772 | 10.5 | 88.3543 | -11.6457 | 2.5 | -0.48912 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T237 | ICPMS21055A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0210015 | 2021-03-26 |  | 108.9635 | 117 | 93.1312 | -6.8688 | 11.3 | -0.71119 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS21077A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0270078 | 2021-03-26 |  | 20.9443 | 22 | 95.2014 | -4.7986 | 2.5 | -0.42228 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21077A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 3580059 | 2021-03-26 |  | 191.6181 | 231 | 82.9516 | -17.0484 | 16.8 | -2.3442 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS21077A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0330077 | 2021-04-21 |  | 25.6107 | 25.7 | 99.6525 | -0.34747 | 2.5 | -0.03572 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS21110A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0400076 | 2021-05-04 |  | 23.1244 | 20.7 | 111.7121 | 11.7121 | 2.5 | 0.96976 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21124A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1030144 | 2021-05-05 |  | 113.3362 | 117 | 96.8685 | -3.1315 | 11.3 | -0.32423 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS21125A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0540055 | 2021-05-11 |  | 19.8262 | 20.7 | 95.7787 | -4.2213 | 2.5 | -0.34952 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21131A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0480038 | 2021-05-12 |  | 24.3189 | 25.7 | 94.6261 | -5.3739 | 2.5 | -0.55244 | 5 | 3415 | Sample was diluted |  |  | µg/L | T231 | ICPMS21131A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0620100 | 2021-05-13 |  | 9.7901 | 10.5 | 93.239 | -6.761 | 2.5 | -0.28396 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T237 | ICPMS21132B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0760035 | 2021-05-14 |  | 9.5111 | 10.5 | 90.5819 | -9.4181 | 2.5 | -0.39556 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T237 | ICPMS21134A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0680063 | 2021-05-14 |  | 229.2204 | 231 | 99.2296 | -0.77039 | 16.8 | -0.10593 | 5 | 3415 | Sample was diluted |  |  | µg/L | T239 | ICPMS21132B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0970048 | 2021-05-21 |  | 21.7515 | 22 | 98.8705 | -1.1295 | 2.5 | -0.0994 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21140A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 0830027 | 2021-05-22 |  | 225.6602 | 231 | 97.6884 | -2.3116 | 16.8 | -0.31785 | 5 | 3415 | Sample was diluted |  |  | µg/L | T239 | ICPMS21140A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1110004 | 2021-06-02 |  | 18.5377 | 22 | 84.2623 | -15.7377 | 2.5 | -1.3849 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21153A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1180038 | 2021-06-16 |  | 125.1753 | 117 | 106.9874 | 6.9874 | 11.3 | 0.72348 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS21167B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1250004 | 2021-06-16 |  | 26.8619 | 25.7 | 104.521 | 4.521 | 2.5 | 0.46476 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS21167B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1320162 | 2021-06-29 |  | 21.4455 | 20.7 | 103.6014 | 3.6014 | 2.5 | 0.2982 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21180B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1380139 | 2021-07-02 |  | 26.604 | 25.7 | 103.5175 | 3.5175 | 2.5 | 0.3616 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS21182B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1460003 | 2021-07-10 |  | 21.2482 | 20.7 | 102.6483 | 2.6483 | 2.5 | 0.21928 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21188B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1540139 | 2021-07-16 |  | 218.4821 | 231 | 94.581 | -5.419 | 16.8 | -0.74511 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS21195A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1600035 | 2021-07-24 |  | 11.1283 | 10.5 | 105.9838 | 5.9838 | 2.5 | 0.25132 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS21200B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1730218 | 2021-08-07 |  | 11.3 | 10.5 | 107.619 | 7.619 | 2.5 | 0.32 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS21218A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1660119 | 2021-08-10 |  | 250.5374 | 231 | 108.4577 | 8.4577 | 16.8 | 1.1629 | 5 | 3415 | Sample was diluted |  |  | µg/L | T239 | ICPMS21215A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1880114 | 2021-08-11 |  | 124.6657 | 117 | 106.5519 | 6.5519 | 11.3 | 0.67838 | 5 | 3415 |  |  |  | µg/L | T229 | ICPMS21222B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 1940229 | 2021-08-16 |  | 19.6381 | 22 | 89.2641 | -10.7359 | 2.5 | -0.94476 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21228B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2080172 | 2021-08-19 |  | 21.4765 | 22 | 97.6205 | -2.3795 | 2.5 | -0.2094 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21230A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2150046 | 2021-08-30 |  | 19.5718 | 20.7 | 94.5498 | -5.4502 | 2.5 | -0.45128 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21242B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2010231 | 2021-09-14 |  | 115.4923 | 117 | 98.7114 | -1.2886 | 11.3 | -0.13342 | 5 | 3415 | Sample was diluted |  |  | µg/L | T229 | ICPMS21252A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2220178 | 2021-09-23 |  | 27.564 | 25.7 | 107.2529 | 7.2529 | 2.5 | 0.7456 | 5 | 3415 |  |  |  | µg/L | T231 | ICPMS21258B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2290121 | 2021-09-24 |  | 22.4576 | 20.7 | 108.4908 | 8.4908 | 2.5 | 0.70304 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS21267A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2360118 | 2021-10-19 |  | 24.9328 | 25.7 | 97.0148 | -2.9852 | 2.5 | -0.30688 | 5 | 3415 | Sample was diluted |  |  | µg/L | T231 | ICPMS21292A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2580265 | 2021-10-29 |  | 224.7569 | 231 | 97.2974 | -2.7026 | 16.8 | -0.37161 | 5 | 3415 | Sample was diluted |  |  | µg/L | T239 | ICPMS21300B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2021 | 2790034 | 2021-11-23 |  | 232.5715 | 231 | 100.6803 | 0.6803 | 16.8 | 0.093542 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS21308B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2870243 | 2021-12-16 |  | 8.3403 | 8.67 | 96.1972 | -3.8028 | 2.5 | -0.13188 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS21349B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2790037 | 2021-12-22 |  | 18.6838 | 22 | 84.9264 | -15.0736 | 2.5 | -1.3265 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS21356A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3000133 | 2022-01-03 |  | 8.3139 | 8.67 | 95.8927 | -4.1073 | 2.5 | -0.14244 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS21356B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3060209 | 2022-01-04 |  | 278.8281 | 275 | 101.392 | 1.392 | 14.6 | 0.2622 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS21356B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3130184 | 2022-01-05 |  | 21.9041 | 20.7 | 105.8169 | 5.8169 | 2.5 | 0.48164 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22005A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3200088 | 2022-01-06 |  | 330 | 275 | 120 | 20 | 14.6 | 3.7671 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS22005A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3410064 | 2022-01-15 |  | 236 | 231 | 102.1645 | 2.1645 | 16.8 | 0.29762 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22014B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3270288 | 2022-01-20 |  | 21 | 20.7 | 101.4493 | 1.4493 | 2.5 | 0.12 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22011B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3480118 | 2022-01-21 |  | 10 | 10.5 | 95.2381 | -4.7619 | 2.5 | -0.2 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22011B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3570003 | 2022-01-28 |  | 260 | 231 | 112.5541 | 12.5541 | 16.8 | 1.7262 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22027C
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 3640134 | 2022-01-29 |  | 12 | 10.5 | 114.2857 | 14.2857 | 2.5 | 0.6 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22027D
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0120120 | 2022-02-17 |  | 22 | 22 | 100 | 0 | 2.5 | 0 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS22047B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0050041 | 2022-02-17 |  | 9 | 8.67 | 103.8062 | 3.8062 | 2.5 | 0.132 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22047B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0250044 | 2022-03-02 |  | 23 | 22 | 104.5455 | 4.5455 | 2.5 | 0.4 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS22060A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0190121 | 2022-03-02 |  | 10 | 8.67 | 115.3403 | 15.3403 | 2.5 | 0.532 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22060A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0380065 | 2022-03-17 |  | 21 | 20.7 | 101.4493 | 1.4493 | 2.5 | 0.12 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22076B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0390071 | 2022-03-18 |  | 302 | 275 | 109.8182 | 9.8182 | 14.6 | 1.8493 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS22076B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0480131 | 2022-03-24 |  | 22 | 20.7 | 106.2802 | 6.2802 | 2.5 | 0.52 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22080A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1090127 | 2022-05-19 |  | 22 | 22 | 100 | 0 | 2.5 | 0 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS22139A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0670123 | 2022-05-19 |  | 275 | 231 | 119.0476 | 19.0476 | 16.8 | 2.619 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22138A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0540113 | 2022-05-19 |  | 328 | 275 | 119.2727 | 19.2727 | 14.6 | 3.6301 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS22138A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0960123 | 2022-05-23 |  | 21 | 22 | 95.4545 | -4.5455 | 2.5 | -0.4 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS22143B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0760093 | 2022-05-24 |  | 11 | 10.5 | 104.7619 | 4.7619 | 2.5 | 0.2 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22143B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0880096 | 2022-05-24 |  | 11 | 10.5 | 104.7619 | 4.7619 | 2.5 | 0.2 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22143B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1170082 | 2022-05-25 |  | 9 | 8.67 | 103.8062 | 3.8062 | 2.5 | 0.132 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22144A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1020024 | 2022-05-26 |  | 10 | 8.67 | 115.3403 | 15.3403 | 2.5 | 0.532 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22146A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1380070 | 2022-06-24 |  | 299 | 275 | 108.7273 | 8.7273 | 14.6 | 1.6438 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS22174A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1230094 | 2022-06-25 |  | 295 | 275 | 107.2727 | 7.2727 | 14.6 | 1.3699 | 5 | 3415 | Sample was diluted |  |  | µg/L | T243 | ICPMS22175A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1460098 | 2022-06-27 |  | 21 | 20.7 | 101.4493 | 1.4493 | 2.5 | 0.12 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22178A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 0820073 | 2022-07-01 |  | 240 | 231 | 103.8961 | 3.8961 | 16.8 | 0.53571 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22181A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1320258 | 2022-07-19 |  | 22 | 20.7 | 106.2802 | 6.2802 | 2.5 | 0.52 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22199A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1580091 | 2022-07-22 |  | 248 | 231 | 107.3593 | 7.3593 | 16.8 | 1.0119 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22202A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1660162 | 2022-07-26 |  | 11 | 10.5 | 104.7619 | 4.7619 | 2.5 | 0.2 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22203B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1790078 | 2022-07-28 |  | 10 | 10.5 | 95.2381 | -4.7619 | 2.5 | -0.2 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22208B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1730126 | 2022-07-28 |  | 239 | 231 | 103.4632 | 3.4632 | 16.8 | 0.47619 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22208B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1870090 | 2022-08-02 |  | 9 | 8.67 | 103.8062 | 3.8062 | 2.5 | 0.132 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22214B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 1930200 | 2022-08-04 |  | 21 | 22 | 95.4545 | -4.5455 | 2.5 | -0.4 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS22215B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2000024 | 2022-08-04 |  | 8 | 8.67 | 92.2722 | -7.7278 | 2.5 | -0.268 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22215B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2080416 | 2022-08-08 |  | 21 | 22 | 95.4545 | -4.5455 | 2.5 | -0.4 | 5 | 3415 |  |  |  | µg/L | T233 | ICPMS22220A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2140174 | 2022-08-09 |  | 22 | 20.7 | 106.2802 | 6.2802 | 2.5 | 0.52 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22220A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2210114 | 2022-08-19 |  | 298 | 275 | 108.3636 | 8.3636 | 14.6 | 1.5753 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS22229B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2280351 | 2022-08-26 |  | 22 | 20.7 | 106.2802 | 6.2802 | 2.5 | 0.52 | 5 | 3415 |  |  |  | µg/L | T235 | ICPMS22238A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2360390 | 2022-09-21 |  | 287 | 275 | 104.3636 | 4.3636 | 14.6 | 0.82192 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS22263B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2570117 | 2022-09-28 |  | 249 | 231 | 107.7922 | 7.7922 | 16.8 | 1.0714 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS22269A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2500494 | 2022-10-05 |  | 11 | 10.5 | 104.7619 | 4.7619 | 2.5 | 0.2 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22277A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2022 | 2630184 | 2022-10-05 |  | 12 | 10.5 | 114.2857 | 14.2857 | 2.5 | 0.6 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS22277A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2770114 | 2022-10-26 |  | 9 | 8.24 | 109.2233 | 9.2233 | 2.5 | 0.304 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS22298B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2860101 | 2022-10-28 |  | 9 | 8.67 | 103.8062 | 3.8062 | 2.5 | 0.132 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS22300A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2920222 | 2022-10-28 |  | 8 | 8.24 | 97.0874 | -2.9126 | 2.5 | -0.096 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS22300A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3610135 | 2023-01-20 |  | 249.7191 | 231 | 108.1035 | 8.1035 | 16.8 | 1.1142 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS23019A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3060049 | 2023-01-20 |  | 293.5614 | 275 | 106.7496 | 6.7496 | 14.6 | 1.2713 | 5 | 3415 | Sample was diluted |  |  | µg/L | T243 | ICPMS23019A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3130109 | 2023-01-20 |  | 144.8726 | 142 | 102.023 | 2.023 | 9.08 | 0.31637 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23019A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3200145 | 2023-01-28 |  | 324.6931 | 275 | 118.0702 | 18.0702 | 14.6 | 3.4036 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS23027A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3400090 | 2023-02-02 |  | 260.0628 | 231 | 112.5813 | 12.5813 | 16.8 | 1.7299 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS23032A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0250122 | 2023-02-02 |  | 8.5139 | 8.67 | 98.1995 | -1.8005 | 2.5 | -0.06244 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS23032A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3330054 | 2023-02-02 |  | 158.1907 | 142 | 111.4019 | 11.4019 | 9.08 | 1.7831 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23032A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 3490121 | 2023-02-09 |  | 9.9077 | 10.5 | 94.359 | -5.641 | 2.5 | -0.23692 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T237 | ICPMS23040A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0040114 | 2023-02-14 |  | 10.2765 | 10.5 | 97.8714 | -2.1286 | 2.5 | -0.0894 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS23045A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0100103 | 2023-02-14 |  | 8.4693 | 8.67 | 97.6851 | -2.3149 | 2.5 | -0.08028 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS23045A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0190114 | 2023-02-16 |  | 8.062 | 8.24 | 97.8398 | -2.1602 | 2.5 | -0.0712 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS23047A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0310087 | 2023-02-17 |  | 8.626 | 8.24 | 104.6845 | 4.6845 | 2.5 | 0.1544 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS23048A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0380151 | 2023-02-18 |  | 160.916 | 142 | 113.3211 | 13.3211 | 9.08 | 2.0833 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23048A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0590092 | 2023-03-16 |  | 141.1845 | 142 | 99.4257 | -0.5743 | 9.08 | -0.089813 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T247 | ICPMS23074A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0740033 | 2023-03-22 |  | 250.4284 | 231 | 108.4106 | 8.4106 | 16.8 | 1.1565 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS23080A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0670102 | 2023-03-25 |  | 11.0369 | 10.5 | 105.1133 | 5.1133 | 2.5 | 0.21476 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T237 | ICPMS23083B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0620097 | 2023-03-25 |  | 312.2401 | 275 | 113.5419 | 13.5419 | 14.6 | 2.5507 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS23083B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0800075 | 2023-03-30 |  | 10.8709 | 10.5 | 103.5324 | 3.5324 | 2.5 | 0.14836 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T237 | ICPMS23089B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0880093 | 2023-04-28 |  | 267.8575 | 231 | 115.9556 | 15.9556 | 16.8 | 2.1939 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS23117A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1020111 | 2023-04-28 |  | 8.0324 | 8.67 | 92.6459 | -7.3541 | 2.5 | -0.25504 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS23117A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0470007 | 2023-04-28 |  | 334.759 | 275 | 121.7305 | 21.7305 | 14.6 | 4.0931 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS23117A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 0940130 | 2023-04-28 |  | 7.5998 | 8.24 | 92.2306 | -7.7694 | 2.5 | -0.25608 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS23117A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1100089 | 2023-05-05 |  | 7.9825 | 8.24 | 96.875 | -3.125 | 2.5 | -0.103 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS23125C
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1240210 | 2023-05-20 |  | 323.5122 | 275 | 117.6408 | 17.6408 | 14.6 | 3.3228 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T243 | ICPMS23139A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1170147 | 2023-05-23 |  | 9.0491 | 8.67 | 104.3725 | 4.3725 | 2.5 | 0.15164 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T241 | ICPMS23142B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1440060 | 2023-06-06 |  | 159.6937 | 142 | 112.4604 | 12.4604 | 9.08 | 1.9486 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23156B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1300167 | 2023-06-10 |  | 165.336 | 142 | 116.4338 | 16.4338 | 9.08 | 2.57 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23160A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1360048 | 2023-06-22 |  | 327.0307 | 275 | 118.9203 | 18.9203 | 14.6 | 3.5637 | 5 | 3415 | Sample was diluted |  |  | µg/L | T243 | ICPMS23171A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1650066 | 2023-06-26 |  | 11.6887 | 10.5 | 111.321 | 11.321 | 2.5 | 0.47548 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T237 | ICPMS23177B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1570041 | 2023-07-18 |  | 248.2179 | 231 | 107.4536 | 7.4536 | 16.8 | 1.0249 | 5 | 3415 |  |  |  | µg/L | T239 | ICPMS23198A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1870310 | 2023-07-18 |  | 11.7419 | 8.67 | 135.4314 | 35.4314 | 2.5 | 1.2288 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T241 | ICPMS23198A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1990166 | 2023-07-28 |  | 9.2972 | 8.67 | 107.2341 | 7.2341 | 2.5 | 0.25088 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria | Below the reporting level but at or above the detection level |  | µg/L | T241 | ICPMS23208B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1940234 | 2023-07-28 |  | 8.3393 | 8.24 | 101.2051 | 1.2051 | 2.5 | 0.03972 | 5 | 3415 | Lab control sample (LCS) recovery outside of range or criteria | Below the reporting level but at or above the detection level |  | µg/L | T245 | ICPMS23208B
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1840086 | 2023-08-01 |  | 11.0239 | 10.5 | 104.9895 | 4.9895 | 2.5 | 0.20956 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS23213A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 1730211 | 2023-08-01 |  | 229.2254 | 231 | 99.2318 | -0.76823 | 16.8 | -0.10563 | 5 | 3415 | Sample was diluted |  |  | µg/L | T239 | ICPMS23213A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2060142 | 2023-08-08 |  | 9.2851 | 8.24 | 112.6833 | 12.6833 | 2.5 | 0.41804 | 5 | 3415 | Below the reporting level but at or above the detection level |  |  | µg/L | T245 | ICPMS23219A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2280173 | 2023-09-08 |  | 314.916 | 275 | 114.5149 | 14.5149 | 14.6 | 2.734 | 5 | 3415 |  |  |  | µg/L | T243 | ICPMS23250A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2490138 | 2023-09-18 |  | 11.3313 | 10.5 | 107.9171 | 7.9171 | 2.5 | 0.33252 | 5 | 3415 |  |  |  | µg/L | T237 | ICPMS23257A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2200127 | 2023-09-27 |  | 157.7458 | 142 | 111.0886 | 11.0886 | 9.08 | 1.7341 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23269A
Boron, UNF, ICP-MS | 01022 | PLM78 | UNF | 2023 | 2350096 | 2023-09-27 |  | 157.3195 | 142 | 110.7884 | 10.7884 | 9.08 | 1.6872 | 5 | 3415 |  |  |  | µg/L | T247 | ICPMS23269A