## IBSP Data
Parameter Description | Parmcode | Methcode | Phase | Water Year | Lab Sample ID | Analysis Date | Remark Code | Lab Value | Expected Conc. | % Recovery | % Bias | 1 QCU | # of QCUs | Report Level | Labcode | Value Qualifier1 | Value Qualifier2 | Value Qualifier3 | Units | Blind Sample ID | Setno
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2760033 | 2018-10-17 |  | 0.19341 | 0.2 | 96.705 | -3.295 | 0.012 | -0.54917 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18289A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2830103 | 2018-10-17 |  | 0.19128 | 0.2 | 95.64 | -4.36 | 0.012 | -0.72667 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18289A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2770143 | 2018-10-17 |  | 0.07123 | 0.0663 | 107.4359 | 7.4359 | 0.005 | 0.986 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18289A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2840113 | 2018-10-17 |  | 0.07 | 0.0663 | 105.5807 | 5.5807 | 0.005 | 0.74 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18289A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2890059 | 2018-10-23 |  | 0.20108 | 0.2 | 100.54 | 0.54 | 0.012 | 0.09 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18295A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2910157 | 2018-10-23 |  | 0.06944 | 0.0663 | 104.736 | 4.736 | 0.005 | 0.628 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18295A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3120138 | 2018-11-14 |  | 0.19106 | 0.2 | 95.53 | -4.47 | 0.012 | -0.745 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18317A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3100175 | 2018-11-14 |  | 0.06841 | 0.0663 | 103.1825 | 3.1825 | 0.005 | 0.422 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18317A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3180069 | 2018-11-26 |  | 0.06321 | 0.0663 | 95.3394 | -4.6606 | 0.005 | -0.618 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18325A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3250021 | 2018-12-11 |  | 0.19193 | 0.2 | 95.965 | -4.035 | 0.012 | -0.6725 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18345A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3310052 | 2018-12-11 |  | 0.19322 | 0.2 | 96.61 | -3.39 | 0.012 | -0.565 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18345A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3400028 | 2018-12-11 |  | 0.19283 | 0.2 | 96.415 | -3.585 | 0.012 | -0.5975 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL18345A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3270029 | 2018-12-11 |  | 0.06585 | 0.0663 | 99.3213 | -0.67873 | 0.005 | -0.09 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18345A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3410002 | 2018-12-11 |  | 0.06498 | 0.0663 | 98.009 | -1.991 | 0.005 | -0.264 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL18345A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 3530020 | 2019-01-30 |  | 0.18596 | 0.2 | 92.98 | -7.02 | 0.012 | -1.17 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N139 | KNLL19029A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0520009 | 2019-03-07 |  | 0.19561 | 0.2 | 97.805 | -2.195 | 0.012 | -0.36583 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL19065A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0520136 | 2019-03-07 |  | 0.06213 | 0.0663 | 93.7104 | -6.2896 | 0.005 | -0.834 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL19065A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0660076 | 2019-03-11 |  | 0.19183 | 0.2 | 95.915 | -4.085 | 0.012 | -0.68083 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL19067A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0640095 | 2019-03-11 |  | 0.06484 | 0.0663 | 97.7979 | -2.2021 | 0.005 | -0.292 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL19067A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0730068 | 2019-03-19 |  | 0.19018 | 0.2 | 95.09 | -4.91 | 0.012 | -0.81833 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL19077A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0720021 | 2019-03-19 |  | 0.06759 | 0.0663 | 101.9457 | 1.9457 | 0.005 | 0.258 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL19077A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0800129 | 2019-03-25 |  | 0.18603 | 0.2 | 93.015 | -6.985 | 0.012 | -1.1642 | 0.01 | 3157 |  |  |  | mg/L as N | N139 | KNLL19084A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0780098 | 2019-03-25 |  | 0.06401 | 0.0663 | 96.546 | -3.454 | 0.005 | -0.458 | 0.01 | 3157 |  |  |  | mg/L as N | N140X | KNLL19084A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0920067 | 2019-04-11 |  | 0.16544 | 0.141 | 117.3333 | 17.3333 | 0.009 | 2.7156 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19101A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0990131 | 2019-04-11 |  | 0.13968 | 0.141 | 99.0638 | -0.93617 | 0.009 | -0.14667 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19101A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 0940210 | 2019-04-11 |  | 0.04288 | 0.046 | 93.2174 | -6.7826 | 0.005 | -0.624 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19101A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1060104 | 2019-04-17 |  | 0.13792 | 0.141 | 97.8156 | -2.1844 | 0.009 | -0.34222 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19107A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1010107 | 2019-04-17 |  | 0.04856 | 0.046 | 105.5652 | 5.5652 | 0.005 | 0.512 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19107A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1080308 | 2019-04-26 |  | 0.04795 | 0.046 | 104.2391 | 4.2391 | 0.005 | 0.39 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19116A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1290100 | 2019-05-14 |  | 0.13701 | 0.141 | 97.1702 | -2.8298 | 0.009 | -0.44333 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19134A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1280119 | 2019-05-14 |  | 0.05107 | 0.046 | 111.0217 | 11.0217 | 0.005 | 1.014 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19134A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1360141 | 2019-05-23 |  | 0.13956 | 0.141 | 98.9787 | -1.0213 | 0.009 | -0.16 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19143A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1340208 | 2019-05-23 |  | 0.04691 | 0.046 | 101.9783 | 1.9783 | 0.005 | 0.182 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19143A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1410228 | 2019-05-23 |  | 0.04885 | 0.046 | 106.1957 | 6.1957 | 0.005 | 0.57 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19143A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1430270 | 2019-05-30 |  | 0.13819 | 0.141 | 98.0071 | -1.9929 | 0.009 | -0.31222 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19149A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1550164 | 2019-06-10 |  | 0.13563 | 0.141 | 96.1915 | -3.8085 | 0.009 | -0.59667 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19161A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1570212 | 2019-06-12 |  | 0.04705 | 0.046 | 102.2826 | 2.2826 | 0.005 | 0.21 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19162A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1620133 | 2019-06-20 |  | 0.13752 | 0.141 | 97.5319 | -2.4681 | 0.009 | -0.38667 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19171A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1690139 | 2019-06-20 |  | 0.13411 | 0.141 | 95.1135 | -4.8865 | 0.009 | -0.76556 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19171A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1640204 | 2019-06-20 |  | 0.0457 | 0.046 | 99.3478 | -0.65217 | 0.005 | -0.06 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19171A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1710259 | 2019-06-27 |  | 0.04278 | 0.046 | 93 | -7 | 0.005 | -0.644 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19178A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1840118 | 2019-07-10 |  | 0.13153 | 0.141 | 93.2837 | -6.7163 | 0.009 | -1.0522 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19190A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1830102 | 2019-07-10 |  | 0.04507 | 0.046 | 97.9783 | -2.0217 | 0.005 | -0.186 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19190A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1900103 | 2019-07-10 |  | 0.04526 | 0.046 | 98.3913 | -1.6087 | 0.005 | -0.148 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19190A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 1930196 | 2019-07-17 |  | 0.13277 | 0.141 | 94.1631 | -5.8369 | 0.009 | -0.91444 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19198A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2000020 | 2019-07-24 |  | 0.04138 | 0.046 | 89.9565 | -10.0435 | 0.005 | -0.924 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19205A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2030081 | 2019-07-29 |  | 0.13341 | 0.141 | 94.617 | -5.383 | 0.009 | -0.84333 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19210A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2180153 | 2019-08-12 |  | 0.13852 | 0.141 | 98.2411 | -1.7589 | 0.009 | -0.27556 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19224A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2200201 | 2019-08-12 |  | 0.04906 | 0.046 | 106.6522 | 6.6522 | 0.005 | 0.612 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19224A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2250167 | 2019-08-20 |  | 0.13607 | 0.141 | 96.5035 | -3.4965 | 0.009 | -0.54778 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19232A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2270235 | 2019-08-20 |  | 0.04858 | 0.046 | 105.6087 | 5.6087 | 0.005 | 0.516 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19232A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2320183 | 2019-08-29 |  | 0.1412 | 0.141 | 100.1418 | 0.14184 | 0.009 | 0.022222 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19241A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2340258 | 2019-08-29 |  | 0.05093 | 0.046 | 110.7174 | 10.7174 | 0.005 | 0.986 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19241A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2470183 | 2019-09-05 |  | 0.04851 | 0.046 | 105.4565 | 5.4565 | 0.005 | 0.502 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19248A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2480209 | 2019-09-13 |  | 0.13745 | 0.141 | 97.4823 | -2.5177 | 0.009 | -0.39444 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19254A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2550079 | 2019-09-13 |  | 0.04393 | 0.046 | 95.5 | -4.5 | 0.005 | -0.414 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19254A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2560160 | 2019-09-26 |  | 0.13771 | 0.141 | 97.6667 | -2.3333 | 0.009 | -0.36556 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19269A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2620332 | 2019-09-26 |  | 0.1385 | 0.141 | 98.227 | -1.773 | 0.009 | -0.27778 | 0.01 | 3157 |  |  |  | mg/L as N | N141 | KNLL19269A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2019 | 2600330 | 2019-09-26 |  | 0.04679 | 0.046 | 101.7174 | 1.7174 | 0.005 | 0.158 | 0.01 | 3157 |  |  |  | mg/L as N | N142X | KNLL19269A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2810116 | 2019-10-09 |  | 0.17686 | 0.18 | 98.2556 | -1.7444 | 0.011 | -0.28545 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19282A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2890224 | 2019-10-22 |  | 0.1705 | 0.18 | 94.7222 | -5.2778 | 0.011 | -0.86364 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19295A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2830132 | 2019-10-22 |  | 0.05648 | 0.0547 | 103.2541 | 3.2541 | 0.005 | 0.356 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19295A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2910181 | 2019-10-22 |  | 0.05543 | 0.0547 | 101.3346 | 1.3346 | 0.005 | 0.146 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19295A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2950099 | 2019-10-25 |  | 0.18178 | 0.18 | 100.9889 | 0.98889 | 0.011 | 0.16182 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19298A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2970177 | 2019-10-25 |  | 0.06 | 0.0547 | 109.6892 | 9.6892 | 0.005 | 1.06 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19298A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3090143 | 2019-11-06 |  | 0.05648 | 0.0547 | 103.2541 | 3.2541 | 0.005 | 0.356 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19310A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3110126 | 2019-11-19 |  | 0.18457 | 0.18 | 102.5389 | 2.5389 | 0.011 | 0.41545 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19323A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3180075 | 2019-11-19 |  | 0.17953 | 0.18 | 99.7389 | -0.26111 | 0.011 | -0.042727 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19323A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3170100 | 2019-11-19 |  | 0.05833 | 0.0547 | 106.6362 | 6.6362 | 0.005 | 0.726 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19323A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3250087 | 2019-11-29 |  | 0.18059 | 0.18 | 100.3278 | 0.32778 | 0.011 | 0.053636 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19333A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3230097 | 2019-11-29 |  | 0.05575 | 0.0547 | 101.9196 | 1.9196 | 0.005 | 0.21 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19333A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3370099 | 2019-12-10 |  | 0.17478 | 0.18 | 97.1 | -2.9 | 0.011 | -0.47455 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19344A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3390132 | 2019-12-10 |  | 0.05424 | 0.0547 | 99.159 | -0.84095 | 0.005 | -0.092 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL19344A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3510020 | 2019-12-18 |  | 0.17898 | 0.18 | 99.4333 | -0.56667 | 0.011 | -0.092727 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL19352A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 3530003 | 2020-01-07 |  | 0.05975 | 0.0547 | 109.2322 | 9.2322 | 0.005 | 1.01 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20007A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0090080 | 2020-01-23 |  | 0.17108 | 0.18 | 95.0444 | -4.9556 | 0.011 | -0.81091 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20022A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0070045 | 2020-01-23 |  | 0.04429 | 0.0547 | 80.9689 | -19.0311 | 0.005 | -2.082 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20022A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0230067 | 2020-02-04 |  | 0.18486 | 0.18 | 102.7 | 2.7 | 0.011 | 0.44182 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20035A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0220112 | 2020-02-04 |  | 0.05769 | 0.0547 | 105.4662 | 5.4662 | 0.005 | 0.598 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20035A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0350139 | 2020-02-19 |  | 0.1783 | 0.18 | 99.0556 | -0.94444 | 0.011 | -0.15455 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20050A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0370097 | 2020-02-19 |  | 0.05764 | 0.0547 | 105.3748 | 5.3748 | 0.005 | 0.588 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20050A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0500056 | 2020-02-25 |  | 0.1724 | 0.18 | 95.7778 | -4.2222 | 0.011 | -0.69091 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20056A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0510087 | 2020-02-25 |  | 0.05546 | 0.0547 | 101.3894 | 1.3894 | 0.005 | 0.152 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20056A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0630049 | 2020-03-05 |  | 0.05943 | 0.0547 | 108.6472 | 8.6472 | 0.005 | 0.946 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20065A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0650092 | 2020-03-18 |  | 0.1688 | 0.18 | 93.7778 | -6.2222 | 0.011 | -1.0182 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20078A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0730162 | 2020-03-18 |  | 0.16971 | 0.18 | 94.2833 | -5.7167 | 0.011 | -0.93545 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20078A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0700105 | 2020-03-18 |  | 0.0552 | 0.0547 | 100.9141 | 0.91408 | 0.005 | 0.1 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20078A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0770113 | 2020-03-18 |  | 0.05428 | 0.0547 | 99.2322 | -0.76782 | 0.005 | -0.084 | 0.01 | 3157 |  |  |  | mg/L as N | N144X | KNLL20078A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0800003 | 2020-03-31 |  | 0.17908 | 0.18 | 99.4889 | -0.51111 | 0.011 | -0.083636 | 0.01 | 3157 |  |  |  | mg/L as N | N143 | KNLL20091A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 0980041 | 2020-04-08 |  | 0.08854 | 0.094 | 94.1915 | -5.8085 | 0.009 | -0.60667 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL20099A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1060057 | 2020-04-16 |  | 0.09368 | 0.094 | 99.6596 | -0.34043 | 0.009 | -0.035556 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL20107A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1000060 | 2020-04-16 |  | 0.04231 | 0.0408 | 103.701 | 3.701 | 0.005 | 0.302 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL20107A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1120051 | 2020-04-22 |  | 0.08654 | 0.094 | 92.0638 | -7.9362 | 0.009 | -0.82889 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL20113A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1070052 | 2020-04-22 |  | 0.0401 | 0.0408 | 98.2843 | -1.7157 | 0.005 | -0.14 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL20113A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1140083 | 2020-05-06 |  | 0.04131 | 0.0408 | 101.25 | 1.25 | 0.005 | 0.102 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL20127A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1260080 | 2020-05-06 |  | 0.04104 | 0.0408 | 100.5882 | 0.58824 | 0.005 | 0.048 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL20127A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1280005 | 2020-05-12 |  | 0.08824 | 0.094 | 93.8723 | -6.1277 | 0.009 | -0.64 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL20133A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1330128 | 2020-05-13 |  | 0.04429 | 0.0408 | 108.5539 | 8.5539 | 0.005 | 0.698 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL20134A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1400017 | 2020-05-28 |  | 0.09265 | 0.094 | 98.5638 | -1.4362 | 0.009 | -0.15 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20149A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1430096 | 2020-05-28 |  | 0.0914 | 0.094 | 97.234 | -2.766 | 0.009 | -0.28889 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20149A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1550010 | 2020-06-10 |  | 0.09041 | 0.094 | 96.1809 | -3.8191 | 0.009 | -0.39889 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20162A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1610194 | 2020-06-10 |  | 0.09304 | 0.094 | 98.9787 | -1.0213 | 0.009 | -0.10667 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20162A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1560173 | 2020-06-10 |  | 0.04305 | 0.0408 | 105.5147 | 5.5147 | 0.005 | 0.45 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20162A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1680151 | 2020-06-24 |  | 0.0892 | 0.094 | 94.8936 | -5.1064 | 0.009 | -0.53333 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20176A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1630142 | 2020-06-24 |  | 0.0422 | 0.0408 | 103.4314 | 3.4314 | 0.005 | 0.28 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20176A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1700256 | 2020-06-24 |  | 0.0424 | 0.0408 | 103.9216 | 3.9216 | 0.005 | 0.32 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20176A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1910117 | 2020-07-17 |  | 0.09285 | 0.094 | 98.7766 | -1.2234 | 0.009 | -0.12778 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20199A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1980095 | 2020-07-17 |  | 0.08721 | 0.094 | 92.7766 | -7.2234 | 0.009 | -0.75444 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20199A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1890184 | 2020-07-17 |  | 0.04191 | 0.0408 | 102.7206 | 2.7206 | 0.005 | 0.222 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20199A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 1960131 | 2020-07-17 |  | 0.0419 | 0.0408 | 102.6961 | 2.6961 | 0.005 | 0.22 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20199A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2040022 | 2020-07-29 |  | 0.04407 | 0.0408 | 108.0147 | 8.0147 | 0.005 | 0.654 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20211A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2170033 | 2020-08-14 |  | 0.08847 | 0.094 | 94.117 | -5.883 | 0.009 | -0.61444 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20222A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2110026 | 2020-08-14 |  | 0.09161 | 0.094 | 97.4574 | -2.5426 | 0.009 | -0.26556 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20222A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2190158 | 2020-08-14 |  | 0.04846 | 0.0408 | 118.7745 | 18.7745 | 0.005 | 1.532 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20222A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2260120 | 2020-08-19 |  | 0.0408 | 0.0408 | 100 | 0 | 0.005 | 0 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20231A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2240094 | 2020-08-26 |  | 0.09107 | 0.094 | 96.883 | -3.117 | 0.009 | -0.32556 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20238A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2320040 | 2020-08-26 |  | 0.09312 | 0.094 | 99.0638 | -0.93617 | 0.009 | -0.097778 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20238A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2370048 | 2020-08-26 |  | 0.04495 | 0.0408 | 110.1716 | 10.1716 | 0.005 | 0.83 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20238A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2540251 | 2020-09-17 |  | 0.08751 | 0.094 | 93.0957 | -6.9043 | 0.009 | -0.72111 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20260A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2530139 | 2020-09-17 |  | 0.04038 | 0.0408 | 98.9706 | -1.0294 | 0.005 | -0.084 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20260A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2590159 | 2020-09-17 |  | 0.03832 | 0.0408 | 93.9216 | -6.0784 | 0.005 | -0.496 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20260A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2610178 | 2020-10-02 |  | 0.09387 | 0.094 | 99.8617 | -0.1383 | 0.009 | -0.014444 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20276A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2670295 | 2020-10-02 |  | 0.04427 | 0.0408 | 108.5049 | 8.5049 | 0.005 | 0.694 | 0.01 | 3157 |  |  |  | mg/L as N | N146X | KNLL_20276A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2020 | 2680085 | 2020-10-08 |  | 0.09026 | 0.094 | 96.0213 | -3.9787 | 0.009 | -0.41556 | 0.01 | 3157 |  |  |  | mg/L as N | N145 | KNLL_20281A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2810030 | 2020-10-08 |  | 0.23216 | 0.237 | 97.9578 | -2.0422 | 0.013 | -0.37231 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_20281A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2880071 | 2020-10-15 |  | 0.23543 | 0.237 | 99.3376 | -0.66245 | 0.013 | -0.12077 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_20288A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2820103 | 2020-10-15 |  | 0.11246 | 0.104 | 108.1346 | 8.1346 | 0.0055667 | 1.5197 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_20288A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2890048 | 2020-10-20 |  | 0.10967 | 0.104 | 105.4519 | 5.4519 | 0.0055667 | 1.0186 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_20294A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2950031 | 2020-11-10 |  | 0.22525 | 0.237 | 95.0422 | -4.9578 | 0.013 | -0.90385 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_20314A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2960118 | 2020-11-10 |  | 0.10201 | 0.104 | 98.0865 | -1.9135 | 0.0055667 | -0.35748 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_20314A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3080109 | 2020-11-23 |  | 0.10382 | 0.104 | 99.8269 | -0.17308 | 0.0055667 | -0.032335 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_20328A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3150102 | 2020-11-23 |  | 0.10952 | 0.104 | 105.3077 | 5.3077 | 0.0055667 | 0.99161 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_20328A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3180089 | 2020-12-14 |  | 0.23026 | 0.237 | 97.1561 | -2.8439 | 0.013 | -0.51846 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N147 | KNLL_20346A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3240159 | 2020-12-14 |  | 0.22819 | 0.237 | 96.2827 | -3.7173 | 0.013 | -0.67769 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_20346A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3350003 | 2020-12-14 |  | 0.22534 | 0.237 | 95.0802 | -4.9198 | 0.013 | -0.89692 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_20346A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3230040 | 2020-12-14 |  | 0.10511 | 0.104 | 101.0673 | 1.0673 | 0.0055667 | 0.1994 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_20346A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3440083 | 2021-01-15 |  | 0.23482 | 0.237 | 99.0802 | -0.91983 | 0.013 | -0.16769 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N147 | KNLL_21004A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3530074 | 2021-01-15 |  | 0.10504 | 0.104 | 101 | 1 | 0.0055667 | 0.18682 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21004A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0150003 | 2021-01-25 |  | 0.23371 | 0.237 | 98.6118 | -1.3882 | 0.013 | -0.25308 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21021A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0050026 | 2021-01-25 |  | 0.1066 | 0.104 | 102.5 | 2.5 | 0.0055667 | 0.46706 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21021A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0210019 | 2021-01-25 |  | 0.10525 | 0.104 | 101.2019 | 1.2019 | 0.0055667 | 0.22455 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21021A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 3660011 | 2021-01-25 |  | 0.10757 | 0.104 | 103.4327 | 3.4327 | 0.0055667 | 0.64131 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21021A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0400078 | 2021-02-18 |  | 0.23589 | 0.237 | 99.5316 | -0.46835 | 0.013 | -0.085385 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21049A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0280001 | 2021-02-18 |  | 0.23302 | 0.237 | 98.3207 | -1.6793 | 0.013 | -0.30615 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21049A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0350086 | 2021-02-18 |  | 0.10602 | 0.104 | 101.9423 | 1.9423 | 0.0055667 | 0.36287 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21049A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0540056 | 2021-02-24 |  | 0.23725 | 0.237 | 100.1055 | 0.10549 | 0.013 | 0.019231 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21055A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0490010 | 2021-02-24 |  | 0.10601 | 0.104 | 101.9327 | 1.9327 | 0.0055667 | 0.36107 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21055A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0620096 | 2021-03-08 |  | 0.10637 | 0.104 | 102.2788 | 2.2788 | 0.0055667 | 0.42574 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21067A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0690090 | 2021-03-18 |  | 0.23349 | 0.237 | 98.519 | -1.481 | 0.013 | -0.27 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21077A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0680065 | 2021-03-18 |  | 0.10656 | 0.104 | 102.4615 | 2.4615 | 0.0055667 | 0.45988 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21077A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0760023 | 2021-03-18 |  | 0.10148 | 0.104 | 97.5769 | -2.4231 | 0.0055667 | -0.45269 | 0.01 | 3157 |  |  |  | mg/L as N | N148X | KNLL_21077A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0770042 | 2021-04-01 |  | 0.23189 | 0.237 | 97.8439 | -2.1561 | 0.013 | -0.39308 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21091A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0840107 | 2021-04-01 |  | 0.2321 | 0.237 | 97.9325 | -2.0675 | 0.013 | -0.37692 | 0.01 | 3157 |  |  |  | mg/L as N | N147 | KNLL_21091A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 0970045 | 2021-04-16 |  | 0.39704 | 0.393 | 101.028 | 1.028 | 0.01965 | 0.2056 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21106A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1030142 | 2021-04-30 |  | 0.38755 | 0.393 | 98.6132 | -1.3868 | 0.01965 | -0.27735 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21120A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1110005 | 2021-04-30 |  | 0.38756 | 0.393 | 98.6158 | -1.3842 | 0.01965 | -0.27684 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21120A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1050216 | 2021-04-30 |  | 0.11414 | 0.114 | 100.1228 | 0.12281 | 0.0059135 | 0.023675 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21120A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1120202 | 2021-04-30 |  | 0.11049 | 0.114 | 96.9211 | -3.0789 | 0.0059135 | -0.59356 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21120A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1190110 | 2021-05-05 |  | 0.11182 | 0.114 | 98.0877 | -1.9123 | 0.0059135 | -0.36865 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21125A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1250007 | 2021-05-14 |  | 0.11747 | 0.114 | 103.0439 | 3.0439 | 0.0059135 | 0.58679 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21134A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1320167 | 2021-05-14 |  | 0.11759 | 0.114 | 103.1491 | 3.1491 | 0.0059135 | 0.60708 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21134A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1330235 | 2021-05-25 |  | 0.39296 | 0.393 | 99.9898 | -0.010178 | 0.01965 | -0.0020356 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21145A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1390135 | 2021-05-25 |  | 0.39414 | 0.393 | 100.2901 | 0.29008 | 0.01965 | 0.058015 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21145A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1380138 | 2021-05-25 |  | 0.11017 | 0.114 | 96.6404 | -3.3596 | 0.0059135 | -0.64767 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21145A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1540138 | 2021-06-09 |  | 0.38232 | 0.393 | 97.2824 | -2.7176 | 0.01965 | -0.54351 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21160A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1530101 | 2021-06-09 |  | 0.38756 | 0.393 | 98.6158 | -1.3842 | 0.01965 | -0.27684 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21160A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1600030 | 2021-06-16 |  | 0.39469 | 0.393 | 100.43 | 0.43003 | 0.01965 | 0.086005 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21167A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1660118 | 2021-06-16 |  | 0.34649 | 0.393 | 88.1654 | -11.8346 | 0.01965 | -2.3669 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21167A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1610182 | 2021-06-16 |  | 0.11638 | 0.114 | 102.0877 | 2.0877 | 0.0059135 | 0.40247 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21167A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1680020 | 2021-06-21 |  | 0.11041 | 0.114 | 96.8509 | -3.1491 | 0.0059135 | -0.60708 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21172A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1750330 | 2021-07-14 |  | 0.11399 | 0.114 | 99.9912 | -0.0087719 | 0.0059135 | -0.001691 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21195A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1960205 | 2021-07-19 |  | 0.381 | 0.393 | 96.9466 | -3.0534 | 0.01965 | -0.61069 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21200A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1880116 | 2021-07-19 |  | 0.11621 | 0.114 | 101.9386 | 1.9386 | 0.0059135 | 0.37372 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21200A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 1940226 | 2021-07-19 |  | 0.11038 | 0.114 | 96.8246 | -3.1754 | 0.0059135 | -0.61216 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21200A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2020247 | 2021-07-23 |  | 0.39671 | 0.393 | 100.944 | 0.94402 | 0.01965 | 0.1888 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21204A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2010232 | 2021-07-23 |  | 0.11645 | 0.114 | 102.1491 | 2.1491 | 0.0059135 | 0.4143 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21204A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2150045 | 2021-08-09 |  | 0.39816 | 0.393 | 101.313 | 1.313 | 0.01965 | 0.2626 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21221A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2100165 | 2021-08-09 |  | 0.39681 | 0.393 | 100.9695 | 0.96947 | 0.01965 | 0.19389 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21221A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2220183 | 2021-08-24 |  | 0.39482 | 0.393 | 100.4631 | 0.4631 | 0.01965 | 0.092621 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21236A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2290125 | 2021-08-24 |  | 0.37642 | 0.393 | 95.7812 | -4.2188 | 0.01965 | -0.84377 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21236A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2240200 | 2021-08-24 |  | 0.11232 | 0.114 | 98.5263 | -1.4737 | 0.0059135 | -0.28409 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21236A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2310207 | 2021-08-30 |  | 0.1063 | 0.114 | 93.2456 | -6.7544 | 0.0059135 | -1.3021 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21242A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2380345 | 2021-09-14 |  | 0.10631 | 0.114 | 93.2544 | -6.7456 | 0.0059135 | -1.3004 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21257A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2510431 | 2021-09-20 |  | 0.11291 | 0.114 | 99.0439 | -0.95614 | 0.0059135 | -0.18432 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21263A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2590009 | 2021-09-27 |  | 0.38401 | 0.393 | 97.7125 | -2.2875 | 0.01965 | -0.45751 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2580268 | 2021-09-27 |  | 0.10677 | 0.114 | 93.6579 | -6.3421 | 0.0059135 | -1.2226 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2660015 | 2021-10-05 |  | 0.40743 | 0.393 | 103.6718 | 3.6718 | 0.01965 | 0.73435 | 0.01 | 3157 |  |  |  | mg/L as N | N149 | KNLL_21278A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2640270 | 2021-10-05 |  | 0.1193 | 0.114 | 104.6491 | 4.6491 | 0.0059135 | 0.89625 | 0.01 | 3157 |  |  |  | mg/L as N | N150X | KNLL_21278A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2021 | 2790019 | 2021-10-13 |  | 1.3897 | 0.393 | 353.6056 | 253.6056 | 0.01965 | 50.7211 | 0.01 | 3157 | Sample was diluted |  |  | mg/L as N | N149 | KNLL_21286A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2790021 | 2021-10-13 |  | 0.8059 | 0.343 | 234.9563 | 134.9563 | 0.019 | 24.3632 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21286A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2870241 | 2021-10-19 |  | 0.33804 | 0.343 | 98.5539 | -1.4461 | 0.019 | -0.26105 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21292A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2940011 | 2021-10-28 |  | 0.31551 | 0.343 | 91.9854 | -8.0146 | 0.019 | -1.4468 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21301A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2940005 | 2021-10-28 |  | 0.20844 | 0.213 | 97.8592 | -2.1408 | 0.01065 | -0.42817 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_21301A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2940111 | 2021-10-28 |  | 0.2027 | 0.213 | 95.1643 | -4.8357 | 0.01065 | -0.96714 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_21301A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3010095 | 2021-11-15 |  | 0.22063 | 0.213 | 103.5822 | 3.5822 | 0.01065 | 0.71643 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_21319A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3060201 | 2021-11-15 |  | 0.22622 | 0.213 | 106.2066 | 6.2066 | 0.01065 | 1.2413 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_21319A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3190110 | 2021-11-23 |  | 0.34627 | 0.343 | 100.9534 | 0.95335 | 0.019 | 0.17211 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21327A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3220099 | 2021-11-23 |  | 0.34 | 0.343 | 99.1254 | -0.87464 | 0.019 | -0.15789 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21327A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3130189 | 2021-11-23 |  | 0.22604 | 0.213 | 106.1221 | 6.1221 | 0.01065 | 1.2244 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_21327A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3200075 | 2021-11-23 |  | 0.22 | 0.213 | 103.2864 | 3.2864 | 0.01065 | 0.65728 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_21327A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3340036 | 2021-12-07 |  | 0.33 | 0.343 | 96.2099 | -3.7901 | 0.019 | -0.68421 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21341A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3410062 | 2021-12-20 |  | 0.34 | 0.343 | 99.1254 | -0.87464 | 0.019 | -0.15789 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_21354A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 3570004 | 2022-01-03 |  | 0.33 | 0.343 | 96.2099 | -3.7901 | 0.019 | -0.68421 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22003A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0050043 | 2022-01-18 |  | 0.21 | 0.213 | 98.5915 | -1.4085 | 0.01065 | -0.28169 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22018A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0130002 | 2022-01-19 |  | 0.34 | 0.343 | 99.1254 | -0.87464 | 0.019 | -0.15789 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22019A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0190123 | 2022-01-20 |  | 0.21 | 0.213 | 98.5915 | -1.4085 | 0.01065 | -0.28169 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22020A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0380063 | 2022-02-09 |  | 0.32 | 0.343 | 93.2945 | -6.7055 | 0.019 | -1.2105 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22039A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0270157 | 2022-02-09 |  | 0.31 | 0.343 | 90.379 | -9.621 | 0.019 | -1.7368 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22039A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0480132 | 2022-02-23 |  | 0.34 | 0.343 | 99.1254 | -0.87464 | 0.019 | -0.15789 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22054A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0470023 | 2022-02-23 |  | 0.21 | 0.213 | 98.5915 | -1.4085 | 0.01065 | -0.28169 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22054A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0560010 | 2022-03-07 |  | 0.22 | 0.213 | 103.2864 | 3.2864 | 0.01065 | 0.65728 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22066A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0690223 | 2022-03-16 |  | 0.33 | 0.343 | 96.2099 | -3.7901 | 0.019 | -0.68421 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22075A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0670120 | 2022-03-16 |  | 0.21 | 0.213 | 98.5915 | -1.4085 | 0.01065 | -0.28169 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22075A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0770096 | 2022-03-22 |  | 0.34 | 0.343 | 99.1254 | -0.87464 | 0.019 | -0.15789 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22081A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0760101 | 2022-03-22 |  | 0.22 | 0.213 | 103.2864 | 3.2864 | 0.01065 | 0.65728 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22081A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0830134 | 2022-03-31 |  | 0.33 | 0.343 | 96.2099 | -3.7901 | 0.019 | -0.68421 | 0.01 | 3157 |  |  |  | mg/L as N | N151 | KNLL_22090A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0880102 | 2022-03-31 |  | 0.21 | 0.213 | 98.5915 | -1.4085 | 0.01065 | -0.28169 | 0.01 | 3157 |  |  |  | mg/L as N | N152X | KNLL_22090A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 0960125 | 2022-04-15 |  | 0.39 | 0.401 | 97.2569 | -2.7431 | 0.02005 | -0.54863 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22105A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1020026 | 2022-04-15 |  | 0.4 | 0.401 | 99.7506 | -0.24938 | 0.02005 | -0.049875 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22105A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1090123 | 2022-05-23 |  | 0.39 | 0.401 | 97.2569 | -2.7431 | 0.02005 | -0.54863 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N153 | KNLL_22143A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1370141 | 2022-06-02 |  | 0.37 | 0.401 | 92.2693 | -7.7307 | 0.02005 | -1.5461 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22153A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1440208 | 2022-06-02 |  | 0.38 | 0.401 | 94.7631 | -5.2369 | 0.02005 | -1.0474 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22153A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1520232 | 2022-06-15 |  | 0.38 | 0.401 | 94.7631 | -5.2369 | 0.02005 | -1.0474 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22166A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1580093 | 2022-07-01 |  | 0.38 | 0.401 | 94.7631 | -5.2369 | 0.02005 | -1.0474 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22182A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1660167 | 2022-07-15 |  | 0.37 | 0.401 | 92.2693 | -7.7307 | 0.02005 | -1.5461 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N153 | KNLL_22196A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1730127 | 2022-07-15 |  | 0.37 | 0.401 | 92.2693 | -7.7307 | 0.02005 | -1.5461 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22196A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 1950117 | 2022-08-11 |  | 0.4 | 0.401 | 99.7506 | -0.24938 | 0.02005 | -0.049875 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22223A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2020023 | 2022-08-11 |  | 0.4 | 0.401 | 99.7506 | -0.24938 | 0.02005 | -0.049875 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22223A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2090179 | 2022-08-30 |  | 0.38 | 0.401 | 94.7631 | -5.2369 | 0.02005 | -1.0474 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N153 | KNLL_22241A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2140178 | 2022-09-02 |  | 0.39 | 0.401 | 97.2569 | -2.7431 | 0.02005 | -0.54863 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N153 | KNLL_22245A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2210119 | 2022-09-02 |  | 0.39 | 0.401 | 97.2569 | -2.7431 | 0.02005 | -0.54863 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22245A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2280356 | 2022-09-15 |  | 0.41 | 0.401 | 102.2444 | 2.2444 | 0.02005 | 0.44888 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N153 | KNLL_22258A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2580149 | 2022-09-29 |  | 0.41 | 0.401 | 102.2444 | 2.2444 | 0.02005 | 0.44888 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22272A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2660046 | 2022-10-17 |  | 0.39 | 0.401 | 97.2569 | -2.7431 | 0.02005 | -0.54863 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22290A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2022 | 2720120 | 2022-10-17 |  | 0.39 | 0.401 | 97.2569 | -2.7431 | 0.02005 | -0.54863 | 0.01 | 3157 |  |  |  | mg/L as N | N153 | KNLL_22290A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2770117 | 2022-10-20 |  | 0.45 | 0.455 | 98.9011 | -1.0989 | 0.026 | -0.19231 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_22293A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2860106 | 2022-10-20 |  | 0.46 | 0.455 | 101.0989 | 1.0989 | 0.026 | 0.19231 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_22293A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2920223 | 2022-11-17 |  | 0.46 | 0.455 | 101.0989 | 1.0989 | 0.026 | 0.19231 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N155 | KNLL_22321A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2910109 | 2022-11-17 |  | 0.45 | 0.451 | 99.7783 | -0.22173 | 0.02255 | -0.044346 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N156X | KNLL_22321A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2930114 | 2022-11-17 |  | 0.44 | 0.451 | 97.561 | -2.439 | 0.02255 | -0.4878 | 0.01 | 3157 | See result laboratory comment |  |  | mg/L as N | N156X | KNLL_22321A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3000168 | 2022-11-21 |  | 0.45 | 0.451 | 99.7783 | -0.22173 | 0.02255 | -0.044346 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_22325A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3060051 | 2022-11-21 |  | 0.45989 | 0.451 | 101.9712 | 1.9712 | 0.02255 | 0.39424 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_22325A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3210271 | 2022-11-22 |  | 0.44982 | 0.455 | 98.8615 | -1.1385 | 0.026 | -0.19923 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_22326A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3130111 | 2022-11-22 |  | 0.44437 | 0.451 | 98.5299 | -1.4701 | 0.02255 | -0.29401 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_22326A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3190027 | 2022-12-02 |  | 0.45173 | 0.455 | 99.2813 | -0.71868 | 0.026 | -0.12577 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_22336A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3340031 | 2022-12-02 |  | 0.46579 | 0.455 | 102.3714 | 2.3714 | 0.026 | 0.415 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_22336A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3200146 | 2022-12-02 |  | 0.46042 | 0.451 | 102.0887 | 2.0887 | 0.02255 | 0.41774 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_22336A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3400092 | 2022-12-21 |  | 0.46389 | 0.455 | 101.9538 | 1.9538 | 0.026 | 0.34192 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_22355A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3540136 | 2023-01-06 |  | 0.47327 | 0.451 | 104.9379 | 4.9379 | 0.02255 | 0.98758 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23006A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 3610136 | 2023-01-24 |  | 0.49206 | 0.455 | 108.1451 | 8.1451 | 0.026 | 1.4254 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N155 | KNLL_23024A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0100104 | 2023-01-24 |  | 0.48772 | 0.451 | 108.1419 | 8.1419 | 0.02255 | 1.6284 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23024A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0050054 | 2023-01-24 |  | 0.48905 | 0.451 | 108.4368 | 8.4368 | 0.02255 | 1.6874 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23024A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0240119 | 2023-02-02 |  | 0.45064 | 0.455 | 99.0418 | -0.95824 | 0.026 | -0.16769 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23033A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0320034 | 2023-02-02 |  | 0.44663 | 0.455 | 98.1604 | -1.8396 | 0.026 | -0.32192 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23033A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0250119 | 2023-02-02 |  | 0.47059 | 0.451 | 104.3437 | 4.3437 | 0.02255 | 0.86874 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23033A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0380148 | 2023-02-09 |  | 0.46833 | 0.455 | 102.9297 | 2.9297 | 0.026 | 0.51269 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23040A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0540081 | 2023-02-24 |  | 0.43862 | 0.451 | 97.255 | -2.745 | 0.02255 | -0.549 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23055A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0590094 | 2023-03-02 |  | 0.48066 | 0.455 | 105.6396 | 5.6396 | 0.026 | 0.98692 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23061A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0660113 | 2023-03-27 |  | 0.46684 | 0.451 | 103.5122 | 3.5122 | 0.02255 | 0.70244 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23086A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0670100 | 2023-03-27 |  | 0.46461 | 0.451 | 103.0177 | 3.0177 | 0.02255 | 0.60355 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23086A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0800082 | 2023-03-27 |  | 0.46028 | 0.451 | 102.0576 | 2.0576 | 0.02255 | 0.41153 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23086A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0750100 | 2023-03-31 |  | 0.44551 | 0.455 | 97.9143 | -2.0857 | 0.026 | -0.365 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23090A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0820083 | 2023-03-31 |  | 0.45321 | 0.455 | 99.6066 | -0.39341 | 0.026 | -0.068846 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23090A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0740034 | 2023-03-31 |  | 0.45225 | 0.451 | 100.2772 | 0.27716 | 0.02255 | 0.055432 | 0.01 | 3157 |  |  |  | mg/L as N | N156X | KNLL_23090A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0890107 | 2023-04-27 |  | 0.47668 | 0.455 | 104.7648 | 4.7648 | 0.026 | 0.83385 | 0.01 | 3157 |  |  |  | mg/L as N | N155 | KNLL_23117A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 0940131 | 2023-05-02 |  | 0.30194 | 0.32 | 94.3563 | -5.6438 | 0.016 | -1.1288 | 0.01 | 3157 |  |  |  | mg/L as N | N157 | KNLL_23122A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1020107 | 2023-05-02 |  | 0.30899 | 0.32 | 96.5594 | -3.4406 | 0.016 | -0.68813 | 0.01 | 3157 |  |  |  | mg/L as N | N157 | KNLL_23122A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1100090 | 2023-05-04 |  | 0.29908 | 0.32 | 93.4625 | -6.5375 | 0.016 | -1.3075 | 0.01 | 3157 |  |  |  | mg/L as N | N157 | KNLL_23124A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1090192 | 2023-05-04 |  | 0.14662 | 0.155 | 94.5935 | -5.4065 | 0.00775 | -1.0813 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23124A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1150153 | 2023-05-04 |  | 0.14356 | 0.155 | 92.6194 | -7.3806 | 0.00775 | -1.4761 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23124A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1230177 | 2023-05-08 |  | 0.14929 | 0.155 | 96.3161 | -3.6839 | 0.00775 | -0.73677 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23128A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1240213 | 2023-05-08 |  | 0.1595 | 0.155 | 102.9032 | 2.9032 | 0.00775 | 0.58065 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23128A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1300165 | 2023-05-23 |  | 0.15933 | 0.155 | 102.7935 | 2.7935 | 0.00775 | 0.55871 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23143A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1310001 | 2023-06-14 |  | 0.31189 | 0.32 | 97.4656 | -2.5344 | 0.016 | -0.50688 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23165A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1380204 | 2023-06-14 |  | 0.34885 | 0.32 | 109.0156 | 9.0156 | 0.016 | 1.8031 | 0.01 | 3157 |  |  |  | mg/L as N | N157 | KNLL_23165A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1360050 | 2023-06-14 |  | 0.15721 | 0.155 | 101.4258 | 1.4258 | 0.00775 | 0.28516 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23165A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1460004 | 2023-06-28 |  | 0.30157 | 0.32 | 94.2406 | -5.7594 | 0.016 | -1.1519 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23179A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1570043 | 2023-09-08 |  | 0.29816 | 0.32 | 93.175 | -6.825 | 0.016 | -1.365 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23250A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1650069 | 2023-09-08 |  | 0.27908 | 0.32 | 87.2125 | -12.7875 | 0.016 | -2.5575 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23250A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1660195 | 2023-09-08 |  | 0.13593 | 0.155 | 87.6968 | -12.3032 | 0.00775 | -2.4606 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23250A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1730213 | 2023-09-21 |  | 0.30158 | 0.32 | 94.2438 | -5.7562 | 0.016 | -1.1512 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23264A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1790032 | 2023-09-21 |  | 0.14189 | 0.155 | 91.5419 | -8.4581 | 0.00775 | -1.6916 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23264A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1980053 | 2023-09-22 |  | 0.30462 | 0.32 | 95.1937 | -4.8063 | 0.016 | -0.96125 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23265A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1860130 | 2023-09-22 |  | 0.14618 | 0.155 | 94.3097 | -5.6903 | 0.00775 | -1.1381 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23265A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1870311 | 2023-09-22 |  | 0.14525 | 0.155 | 93.7097 | -6.2903 | 0.00775 | -1.2581 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23265A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1940236 | 2023-09-22 |  | 0.15027 | 0.155 | 96.9484 | -3.0516 | 0.00775 | -0.61032 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23265A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2020028 | 2023-09-26 |  | 0.30424 | 0.32 | 95.075 | -4.925 | 0.016 | -0.985 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23269A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2080233 | 2023-09-26 |  | 0.30717 | 0.32 | 95.9906 | -4.0094 | 0.016 | -0.80188 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23269A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 1990168 | 2023-09-26 |  | 0.1462 | 0.155 | 94.3226 | -5.6774 | 0.00775 | -1.1355 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23269A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2200131 | 2023-09-27 |  | 0.29591 | 0.32 | 92.4719 | -7.5281 | 0.016 | -1.5056 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2280170 | 2023-09-27 |  | 0.30158 | 0.32 | 94.2438 | -5.7562 | 0.016 | -1.1512 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2350098 | 2023-09-27 |  | 0.29552 | 0.32 | 92.35 | -7.65 | 0.016 | -1.53 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N157 | KNLL_23270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2290215 | 2023-09-27 |  | 0.35642 | 0.155 | 229.9484 | 129.9484 | 0.00775 | 25.9897 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2360089 | 2023-09-27 |  | 0.16162 | 0.155 | 104.271 | 4.271 | 0.00775 | 0.85419 | 0.01 | 3157 | Holding time exceeded | See result laboratory comment |  | mg/L as N | N158X | KNLL_23270A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2570187 | 2023-09-28 |  | 0.30072 | 0.32 | 93.975 | -6.025 | 0.016 | -1.205 | 0.01 | 3157 |  |  |  | mg/L as N | N157 | KNLL_23271A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2490135 | 2023-09-28 |  | 0.14945 | 0.155 | 96.4194 | -3.5806 | 0.00775 | -0.71613 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23271A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2560091 | 2023-09-28 |  | 0.13991 | 0.155 | 90.2645 | -9.7355 | 0.00775 | -1.9471 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23271A
Nitrate plus Nitrite (as N), FIL, DA, LL | 00631 | RED02 | FIL | 2023 | 2480100 | 2023-09-28 |  | 0.16531 | 0.155 | 106.6516 | 6.6516 | 0.00775 | 1.3303 | 0.01 | 3157 |  |  |  | mg/L as N | N158X | KNLL_23271A