## IBSP Data
Parameter Description | Parmcode | Methcode | Phase | Water Year | Lab Sample ID | Analysis Date | Remark Code | Lab Value | Expected Conc. | % Recov | % Bias | 1 QCU | # of QCUs | Report Level | Labcode | Value Qualifier1 | Value Qualifier2 | Value Qualifier3 | Units | Blind Sample ID | Setno
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2019 | 2390203 | 2019-10-01 |  | 187.4741 | 185 | 101.3374 | 1.3374 | 11 | 0.22492 |  | 3309 | Sample was diluted |  |  | ug/L | T231 | ICPMS19274A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2019 | 2470180 | 2019-10-02 |  | 10.3051 | 10.1 | 102.0307 | 2.0307 | 1 | 0.2051 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS19274A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2019 | 2550075 | 2019-10-09 |  | 200.4495 | 205 | 97.7802 | -2.2198 | 11.9 | -0.38239 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS19282B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2019 | 2600328 | 2019-10-09 |  | 11 | 10.1 | 108.9109 | 8.9109 | 1 | 0.9 | 2 | 3309 | Sample was diluted |  |  | µg/L | T225 | ICPMS19282B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2950100 | 2019-12-16 |  | 19.6792 | 19.7 | 99.8944 | -0.10558 | 1.85 | -0.011243 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T233 | ICPMS19347B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2019 | 2680149 | 2019-12-18 |  | 212.1664 | 205 | 103.4958 | 3.4958 | 11.9 | 0.60222 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS19350A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2810114 | 2019-12-19 |  | 18.8356 | 19.7 | 95.6122 | -4.3878 | 1.85 | -0.46724 | 2 | 3309 | Value verified by rerun, same method |  |  | µg/L | T233 | ICPMS19352B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2890221 | 2019-12-19 |  | 233.9083 | 230 | 101.6993 | 1.6993 | 12.3 | 0.31775 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS19352B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3170095 | 2019-12-30 |  | 26.3325 | 24.8 | 106.1794 | 6.1794 | 1.54 | 0.99513 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T235 | ICPMS19364A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3310060 | 2019-12-30 |  | 25.6426 | 24.8 | 103.3976 | 3.3976 | 1.54 | 0.54714 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T235 | ICPMS19364A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3030027 | 2020-01-09 |  | 234.0941 | 230 | 101.78 | 1.78 | 12.3 | 0.33285 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS20006A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3090147 | 2020-01-09 |  | 184.1496 | 185 | 99.5403 | -0.45968 | 11 | -0.077309 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS20006A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3230099 | 2020-01-09 |  | 184.7011 | 185 | 99.8384 | -0.16157 | 11 | -0.027173 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS20006A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3600045 | 2020-01-21 |  | 9.6831 | 10.1 | 95.8723 | -4.1277 | 1 | -0.4169 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T225 | ICPMS20021A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0070044 | 2020-01-30 |  | 250.3776 | 230 | 108.8598 | 8.8598 | 12.3 | 1.6567 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS20030A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3440294 | 2020-01-30 |  | 10.0747 | 10.1 | 99.7495 | -0.2505 | 1 | -0.0253 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS20030A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3370103 | 2020-01-30 |  | 218.3702 | 205 | 106.522 | 6.522 | 11.9 | 1.1235 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20030A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 3510023 | 2020-01-30 |  | 212.1607 | 205 | 103.493 | 3.493 | 11.9 | 0.60174 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20030A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0150017 | 2020-02-03 |  | 19.3508 | 19.7 | 98.2274 | -1.7726 | 1.85 | -0.18876 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20031A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0280070 | 2020-02-03 |  | 19.1577 | 19.7 | 97.2472 | -2.7528 | 1.85 | -0.29314 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20031A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0220109 | 2020-02-05 |  | 220.9587 | 230 | 96.069 | -3.931 | 12.3 | -0.73507 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS20036A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0350135 | 2020-02-12 |  | 23.2512 | 24.8 | 93.7548 | -6.2452 | 1.54 | -1.0057 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T235 | ICPMS20043A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0500059 | 2020-03-04 |  | 22.9697 | 24.8 | 92.6198 | -7.3802 | 1.54 | -1.1885 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T235 | ICPMS20064A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0630051 | 2020-03-18 |  | 10.2921 | 10.1 | 101.902 | 1.902 | 1 | 0.1921 | 2 | 3309 | See result laboratory comment |  |  | µg/L | T225 | ICPMS20076A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0420076 | 2020-04-20 |  | 179.974 | 185 | 97.2832 | -2.7168 | 11 | -0.45691 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS20111B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0560053 | 2020-04-20 |  | 181.4791 | 185 | 98.0968 | -1.9032 | 11 | -0.32008 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS20111B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0980043 | 2020-04-23 |  | 18.425 | 19.7 | 93.5279 | -6.4721 | 1.85 | -0.68919 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20114A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1120049 | 2020-05-04 |  | 19.8018 | 19.7 | 100.5168 | 0.51675 | 1.85 | 0.055027 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20125B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0700103 | 2020-05-04 |  | 206.2646 | 205 | 100.6169 | 0.61688 | 11.9 | 0.10627 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20125B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1060059 | 2020-06-02 |  | 250.7118 | 230 | 109.0051 | 9.0051 | 12.3 | 1.6839 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS20153B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0840029 | 2020-06-02 |  | 213.4884 | 205 | 104.1407 | 4.1407 | 11.9 | 0.71331 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20153B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 0770114 | 2020-06-03 |  | 10.3366 | 10.1 | 102.3426 | 2.3426 | 1 | 0.2366 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS20153B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1330126 | 2020-06-12 |  | 26.4469 | 24.8 | 106.6407 | 6.6407 | 1.54 | 1.0694 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS20163B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1610191 | 2020-07-01 |  | 10.6641 | 10.1 | 105.5851 | 5.5851 | 1 | 0.5641 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS20182B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1760101 | 2020-08-28 |  | 9.6361 | 10.1 | 95.4069 | -4.5931 | 1 | -0.4639 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS20240A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1550007 | 2020-08-28 |  | 201.9378 | 205 | 98.5062 | -1.4938 | 11.9 | -0.25733 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20240A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1960126 | 2020-09-03 |  | 18.5294 | 19.7 | 94.0579 | -5.9421 | 1.85 | -0.63276 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20247B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2170035 | 2020-09-17 |  | 25.9832 | 24.8 | 104.771 | 4.771 | 1.54 | 0.76831 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS20260A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2120199 | 2020-09-17 |  | 20.5938 | 19.7 | 104.5371 | 4.5371 | 1.85 | 0.48314 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20260A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2040025 | 2020-09-21 |  | 236.3205 | 230 | 102.748 | 2.748 | 12.3 | 0.51386 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS20265B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1680149 | 2020-09-21 |  | 207.697 | 205 | 101.3156 | 1.3156 | 11.9 | 0.22664 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20265B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2320037 | 2020-09-24 |  | 24.3813 | 24.8 | 98.3117 | -1.6883 | 1.54 | -0.27188 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS20276A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2530137 | 2020-10-06 |  | 12.6073 | 10.1 | 124.8248 | 24.8248 | 1 | 2.5073 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS20279B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2670298 | 2020-10-09 |  | 10.7361 | 10.1 | 106.298 | 6.298 | 1 | 0.6361 | 2 | 3309 |  |  |  | µg/L | T225 | ICPMS20281B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 1890182 | 2020-10-14 |  | 244.0228 | 230 | 106.0969 | 6.0969 | 12.3 | 1.1401 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS20283A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2810032 | 2020-10-22 |  | 19.6418 | 19.7 | 99.7046 | -0.29543 | 1.85 | -0.031459 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS20296B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2590164 | 2020-12-29 |  | 207.8648 | 205 | 101.3975 | 1.3975 | 11.9 | 0.24074 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20350A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2740056 | 2020-12-29 |  | 208.0574 | 205 | 101.4914 | 1.4914 | 11.9 | 0.25692 | 2 | 3309 | Sample was diluted |  |  | µg/L | T227 | ICPMS20350A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2240092 | 2021-01-11 |  | 193.7941 | 185 | 104.7536 | 4.7536 | 11 | 0.79946 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS20353A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2020 | 2390027 | 2021-01-12 |  | 194.0582 | 185 | 104.8963 | 4.8963 | 11 | 0.82347 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS20353A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2950028 | 2021-01-13 |  | 19.0875 | 19.7 | 96.8909 | -3.1091 | 1.85 | -0.33108 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21013B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2880069 | 2021-01-15 |  | 237.8446 | 230 | 103.4107 | 3.4107 | 12.3 | 0.63777 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21015A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3150100 | 2021-01-15 |  | 26.1815 | 24.8 | 105.5706 | 5.5706 | 1.54 | 0.89708 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21015A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3010155 | 2021-01-27 |  | 243.3793 | 230 | 105.8171 | 5.8171 | 12.3 | 1.0877 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21026B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3080107 | 2021-01-27 |  | 181.203 | 185 | 97.9476 | -2.0524 | 11 | -0.34518 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21026B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3440081 | 2021-01-29 |  | 12.5765 | 13.2 | 95.2765 | -4.7235 | 1.25 | -0.4988 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS21028B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3300047 | 2021-02-02 |  | 27.2171 | 24.8 | 109.7464 | 9.7464 | 1.54 | 1.5695 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21033A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3230039 | 2021-03-04 |  | 190.9073 | 185 | 103.1931 | 3.1931 | 11 | 0.53703 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21055A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0120036 | 2021-03-04 |  | 19.9173 | 19.7 | 101.103 | 1.103 | 1.85 | 0.11746 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21055A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3500073 | 2021-03-04 |  | 276.3586 | 266 | 103.8942 | 3.8942 | 13.3 | 0.77884 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21055A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3580059 | 2021-03-04 |  | 12.4794 | 13.2 | 94.5409 | -5.4591 | 1.25 | -0.57648 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS21055A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0270078 | 2021-03-26 |  | 19.1009 | 19.7 | 96.9589 | -3.0411 | 1.85 | -0.32384 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21077A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0050024 | 2021-03-26 |  | 233.5268 | 230 | 101.5334 | 1.5334 | 12.3 | 0.28673 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21077A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 3640015 | 2021-03-26 |  | 276.554 | 266 | 103.9677 | 3.9677 | 13.3 | 0.79353 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21077A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0400076 | 2021-05-04 |  | 25.2873 | 24.8 | 101.9649 | 1.9649 | 1.54 | 0.31643 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21124A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0330077 | 2021-05-07 |  | 190.7725 | 185 | 103.1203 | 3.1203 | 11 | 0.52477 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21126A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0210015 | 2021-05-08 |  | 243.68 | 230 | 105.9478 | 5.9478 | 12.3 | 1.1122 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21126A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0540055 | 2021-05-11 |  | 25.4697 | 24.8 | 102.7004 | 2.7004 | 1.54 | 0.43487 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21131A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0480038 | 2021-05-12 |  | 189.5587 | 185 | 102.4642 | 2.4642 | 11 | 0.41443 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21131A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0680063 | 2021-05-14 |  | 13.313 | 13.2 | 100.8561 | 0.85606 | 1.25 | 0.0904 | 2 | 3309 | Sample was diluted |  |  | µg/L | T239 | ICPMS21132B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0830027 | 2021-05-15 |  | 13.7928 | 13.2 | 104.4909 | 4.4909 | 1.25 | 0.47424 | 2 | 3309 | Sample was diluted |  |  | µg/L | T239 | ICPMS21134A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0760035 | 2021-05-19 |  | 251.569 | 266 | 94.5748 | -5.4252 | 13.3 | -1.085 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21140A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0970048 | 2021-05-21 |  | 17.8842 | 19.7 | 90.7827 | -9.2173 | 1.85 | -0.98151 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21140A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 0620100 | 2021-05-21 |  | 250.3833 | 266 | 94.1291 | -5.8709 | 13.3 | -1.1742 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21140A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1030144 | 2021-05-26 |  | 239.687 | 230 | 104.2117 | 4.2117 | 12.3 | 0.78756 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21145B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1110004 | 2021-06-02 |  | 18.5416 | 19.7 | 94.1198 | -5.8802 | 1.85 | -0.62616 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21153A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1250004 | 2021-06-23 |  | 189.1324 | 185 | 102.2337 | 2.2337 | 11 | 0.37567 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21173A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1180038 | 2021-06-24 |  | 233.1004 | 230 | 101.348 | 1.348 | 12.3 | 0.25207 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21173A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1320162 | 2021-06-29 |  | 25.2941 | 24.8 | 101.9923 | 1.9923 | 1.54 | 0.32084 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21180B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1460003 | 2021-07-24 |  | 24.2692 | 24.8 | 97.8597 | -2.1403 | 1.54 | -0.34468 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21200B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1380139 | 2021-07-30 |  | 162.5449 | 185 | 87.8621 | -12.1379 | 11 | -2.0414 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21207B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1540139 | 2021-07-31 |  | 13.2069 | 13.2 | 100.0523 | 0.052273 | 1.25 | 0.00552 | 2 | 3309 | Sample was diluted |  |  | µg/L | T239 | ICPMS21207B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1660119 | 2021-08-07 |  | 12.5326 | 13.2 | 94.9439 | -5.0561 | 1.25 | -0.53392 | 2 | 3309 | Sample was diluted |  |  | µg/L | T239 | ICPMS21218A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1600035 | 2021-08-13 |  | 254.4621 | 266 | 95.6624 | -4.3376 | 13.3 | -0.86751 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21225A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1940229 | 2021-08-16 |  | 18.6468 | 19.7 | 94.6538 | -5.3462 | 1.85 | -0.5693 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21228B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2080172 | 2021-08-19 |  | 19.9624 | 19.7 | 101.332 | 1.332 | 1.85 | 0.14184 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21230A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2150046 | 2021-08-30 |  | 26.4652 | 24.8 | 106.7145 | 6.7145 | 1.54 | 1.0813 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21242B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1880114 | 2021-09-01 |  | 238.8883 | 230 | 103.8645 | 3.8645 | 12.3 | 0.72263 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21232A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 1730218 | 2021-09-01 |  | 272.217 | 266 | 102.3372 | 2.3372 | 13.3 | 0.46744 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21232A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2010231 | 2021-09-02 |  | 241.3113 | 230 | 104.918 | 4.918 | 12.3 | 0.91962 | 2 | 3309 | Sample was diluted |  |  | µg/L | T229 | ICPMS21244A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2290121 | 2021-09-24 |  | 25.958 | 24.8 | 104.6694 | 4.6694 | 1.54 | 0.75195 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS21267A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2220178 | 2021-10-02 |  | 188.6442 | 185 | 101.9698 | 1.9698 | 11 | 0.33129 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21273A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2360118 | 2021-10-12 |  | 192.1352 | 185 | 103.8569 | 3.8569 | 11 | 0.64865 | 2 | 3309 | Sample was diluted |  |  | µg/L | T231 | ICPMS21285B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2580265 | 2021-10-20 |  | 13.559 | 13.2 | 102.7197 | 2.7197 | 1.25 | 0.2872 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS21292A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2510433 | 2021-11-02 |  | 262.7045 | 266 | 98.7611 | -1.2389 | 13.3 | -0.24778 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21295B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2640271 | 2021-11-02 |  | 265.2192 | 266 | 99.7065 | -0.29353 | 13.3 | -0.058707 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS21295B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2021 | 2790034 | 2021-11-23 |  | 14.1532 | 13.2 | 107.2212 | 7.2212 | 1.25 | 0.76256 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS21308B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2790037 | 2021-12-14 |  | 20.1834 | 19.7 | 102.4538 | 2.4538 | 1.85 | 0.2613 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS21340B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2870243 | 2021-12-22 |  | 199.3738 | 193 | 103.3025 | 3.3025 | 10.4 | 0.61287 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS21356A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3130184 | 2022-01-05 |  | 25.5727 | 24.8 | 103.1157 | 3.1157 | 1.54 | 0.50175 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22005A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3200088 | 2022-01-15 |  | 453.4634 | 446 | 101.6734 | 1.6734 | 25.6 | 0.29154 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22014B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3410064 | 2022-01-15 |  | 12.9523 | 13.2 | 98.1235 | -1.8765 | 1.25 | -0.19816 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22014B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3000133 | 2022-01-20 |  | 204.632 | 193 | 106.0269 | 6.0269 | 10.4 | 1.1185 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22011B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3270288 | 2022-01-20 |  | 26.4251 | 24.8 | 106.5528 | 6.5528 | 1.54 | 1.0553 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22011B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3570003 | 2022-01-28 |  | 12.8986 | 13.2 | 97.7167 | -2.2833 | 1.25 | -0.24112 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22027C
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0120120 | 2022-02-17 |  | 20.195 | 19.7 | 102.5127 | 2.5127 | 1.85 | 0.26757 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS22047B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0250044 | 2022-03-02 |  | 20.0997 | 19.7 | 102.0289 | 2.0289 | 1.85 | 0.21605 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS22060A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0190121 | 2022-03-02 |  | 196.9494 | 193 | 102.0463 | 2.0463 | 10.4 | 0.37975 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22061B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3640134 | 2022-03-02 |  | 258.8331 | 266 | 97.3057 | -2.6943 | 13.3 | -0.53886 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22060A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3060209 | 2022-03-03 |  | 469.4601 | 446 | 105.2601 | 5.2601 | 25.6 | 0.91641 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22061B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0050041 | 2022-03-11 |  | 199.7819 | 193 | 103.5139 | 3.5139 | 10.4 | 0.65211 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22068B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0380065 | 2022-03-17 |  | 23.0085 | 24.8 | 92.7762 | -7.2238 | 1.54 | -1.1633 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22076B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0480131 | 2022-03-24 |  | 25.0898 | 24.8 | 101.1685 | 1.1685 | 1.54 | 0.18818 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22080A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 3480118 | 2022-03-31 |  | 276.3 | 266 | 103.8722 | 3.8722 | 13.3 | 0.77444 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22095B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0670123 | 2022-04-20 |  | 13.7428 | 13.2 | 104.1121 | 4.1121 | 1.25 | 0.43424 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22088B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0960123 | 2022-05-16 |  | 20.9107 | 19.7 | 106.1457 | 6.1457 | 1.85 | 0.65443 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS22136A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1090127 | 2022-05-19 |  | 19.5416 | 19.7 | 99.1959 | -0.80406 | 1.85 | -0.085622 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS22139A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0820073 | 2022-05-24 |  | 13.7718 | 13.2 | 104.3318 | 4.3318 | 1.25 | 0.45744 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22143B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1020024 | 2022-06-24 |  | 196.2402 | 193 | 101.6789 | 1.6789 | 10.4 | 0.31156 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22175A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1170082 | 2022-06-24 |  | 197.1754 | 193 | 102.1634 | 2.1634 | 10.4 | 0.40148 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22175A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1230094 | 2022-06-25 |  | 468.314 | 446 | 105.0031 | 5.0031 | 25.6 | 0.87164 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22175A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0390071 | 2022-06-30 |  | 482.8307 | 446 | 108.258 | 8.258 | 25.6 | 1.4387 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22181A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0760093 | 2022-06-30 |  | 264.0533 | 266 | 99.2682 | -0.73184 | 13.3 | -0.14637 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22181A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0880096 | 2022-06-30 |  | 264.5515 | 266 | 99.4555 | -0.54455 | 13.3 | -0.10891 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22181A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1320258 | 2022-07-19 |  | 28.3074 | 24.8 | 114.1427 | 14.1427 | 1.54 | 2.2775 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22199A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1380070 | 2022-07-20 |  | 446.6414 | 446 | 100.1438 | 0.14381 | 25.6 | 0.025055 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22199A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 0540113 | 2022-07-20 |  | 460.939 | 446 | 103.3496 | 3.3496 | 25.6 | 0.58355 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22199A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1460098 | 2022-07-22 |  | 25.0719 | 24.8 | 101.0964 | 1.0964 | 1.54 | 0.17656 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22202A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1580091 | 2022-07-22 |  | 12.598 | 13.2 | 95.4394 | -4.5606 | 1.25 | -0.4816 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22202A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1730126 | 2022-07-28 |  | 13.2248 | 13.2 | 100.1879 | 0.18788 | 1.25 | 0.01984 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22208B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1790078 | 2022-08-03 |  | 260.0063 | 266 | 97.7467 | -2.2533 | 13.3 | -0.45065 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22214B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1930200 | 2022-08-04 |  | 20.1847 | 19.7 | 102.4604 | 2.4604 | 1.85 | 0.262 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS22215B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2080416 | 2022-08-08 |  | 18.6949 | 19.7 | 94.898 | -5.102 | 1.85 | -0.5433 | 2 | 3309 |  |  |  | µg/L | T233 | ICPMS22220A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2140174 | 2022-08-09 |  | 24.9798 | 24.8 | 100.725 | 0.725 | 1.54 | 0.11675 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22220A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1870090 | 2022-08-15 |  | 198.6741 | 193 | 102.9399 | 2.9399 | 10.4 | 0.54559 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22223B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2000024 | 2022-08-16 |  | 188.4312 | 193 | 97.6327 | -2.3673 | 10.4 | -0.43931 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22228A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 1660162 | 2022-08-22 |  | 255.9852 | 266 | 96.235 | -3.765 | 13.3 | -0.75299 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22234A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2280351 | 2022-08-26 |  | 26.3777 | 24.8 | 106.3617 | 6.3617 | 1.54 | 1.0245 | 2 | 3309 |  |  |  | µg/L | T235 | ICPMS22238A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2210114 | 2022-08-27 |  | 455.778 | 446 | 102.1924 | 2.1924 | 25.6 | 0.38195 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22238A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2570117 | 2022-09-28 |  | 13.738 | 13.2 | 104.0758 | 4.0758 | 1.25 | 0.4304 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22269A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2360390 | 2022-09-29 |  | 474.7619 | 446 | 106.4489 | 6.4489 | 25.6 | 1.1235 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS22271B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2710051 | 2022-10-26 |  | 13.5308 | 13.2 | 102.5061 | 2.5061 | 1.25 | 0.26464 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS22298B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2500494 | 2022-10-26 |  | 262.9364 | 266 | 98.8483 | -1.1517 | 13.3 | -0.23035 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22298B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2022 | 2630184 | 2022-10-26 |  | 267.9696 | 266 | 100.7405 | 0.74045 | 13.3 | 0.14809 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS22298B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2770114 | 2022-11-01 |  | 163.6239 | 167 | 97.9784 | -2.0216 | 8.35 | -0.40432 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS22305B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2920222 | 2022-11-01 |  | 164.5385 | 167 | 98.526 | -1.474 | 8.35 | -0.29479 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS22305B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2860101 | 2022-11-01 |  | 182.8512 | 193 | 94.7416 | -5.2584 | 10.4 | -0.97585 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22305B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2990136 | 2022-11-08 |  | 185.1544 | 193 | 95.9349 | -4.0651 | 10.4 | -0.75438 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS22312B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3130109 | 2022-11-16 |  | 11.6136 | 10.2 | 113.8588 | 13.8588 | 1.36 | 1.0394 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS22320A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3610135 | 2023-01-09 |  | 13.4317 | 13.2 | 101.7553 | 1.7553 | 1.25 | 0.18536 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23009A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3060049 | 2023-01-20 |  | 465.9809 | 446 | 104.48 | 4.48 | 25.6 | 0.7805 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23019A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3330054 | 2023-02-02 |  | 11.5813 | 10.2 | 113.5422 | 13.5422 | 1.36 | 1.0157 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23032A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0380151 | 2023-02-18 |  | 11.5739 | 10.2 | 113.4696 | 13.4696 | 1.36 | 1.0102 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23048A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0250122 | 2023-02-21 |  | 193.0572 | 193 | 100.0296 | 0.029637 | 10.4 | 0.0055 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS23052A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3400090 | 2023-02-21 |  | 14.3991 | 13.2 | 109.0841 | 9.0841 | 1.25 | 0.95928 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23052A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0590092 | 2023-03-16 |  | 11.5812 | 10.2 | 113.5412 | 13.5412 | 1.36 | 1.0156 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23074A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3200145 | 2023-03-22 |  | 437.9195 | 446 | 98.1882 | -1.8118 | 25.6 | -0.31564 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23080A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0740033 | 2023-03-22 |  | 14.7648 | 13.2 | 111.8545 | 11.8545 | 1.25 | 1.2518 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23080A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 3490121 | 2023-03-22 |  | 256.8738 | 266 | 96.5691 | -3.4309 | 13.3 | -0.68618 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23080A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0040114 | 2023-03-22 |  | 261.9097 | 266 | 98.4623 | -1.5377 | 13.3 | -0.30754 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23080A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0190114 | 2023-03-24 |  | 173.6917 | 167 | 104.007 | 4.007 | 8.35 | 0.8014 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS23083B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0100103 | 2023-03-24 |  | 199.5819 | 193 | 103.4103 | 3.4103 | 10.4 | 0.63287 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS23083B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0470007 | 2023-03-29 |  | 465.4878 | 446 | 104.3695 | 4.3695 | 25.6 | 0.76124 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23088A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0310087 | 2023-03-29 |  | 169.0263 | 167 | 101.2134 | 1.2134 | 8.35 | 0.24267 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS23088A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0620097 | 2023-03-31 |  | 476.6271 | 446 | 106.8671 | 6.8671 | 25.6 | 1.1964 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23089B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0940130 | 2023-05-03 |  | 158.6456 | 167 | 94.9974 | -5.0026 | 8.35 | -1.0005 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS23123B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0670102 | 2023-05-03 |  | 247.0869 | 266 | 92.8898 | -7.1102 | 13.3 | -1.422 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23123B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0800075 | 2023-05-03 |  | 244.4335 | 266 | 91.8923 | -8.1077 | 13.3 | -1.6215 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23123B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 0880093 | 2023-05-04 |  | 12.34 | 13.2 | 93.4848 | -6.5152 | 1.25 | -0.688 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23123B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1300167 | 2023-05-24 |  | 10.0458 | 10.2 | 98.4882 | -1.5118 | 1.36 | -0.11338 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23142B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1100089 | 2023-05-26 |  | 169.3265 | 167 | 101.3931 | 1.3931 | 8.35 | 0.27862 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS23145B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1440060 | 2023-06-06 |  | 10.7975 | 10.2 | 105.8578 | 5.8578 | 1.36 | 0.43934 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23156B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1170147 | 2023-06-10 |  | 194.7408 | 193 | 100.902 | 0.90197 | 10.4 | 0.16738 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS23160A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1020111 | 2023-06-14 |  | 187.3515 | 193 | 97.0733 | -2.9267 | 10.4 | -0.54313 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS23164A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1360048 | 2023-06-15 |  | 481.7921 | 446 | 108.0251 | 8.0251 | 25.6 | 1.3981 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23164A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1570041 | 2023-06-15 |  | 13.0692 | 13.2 | 99.0091 | -0.99091 | 1.25 | -0.10464 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23164A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1240210 | 2023-06-16 |  | 486.1891 | 446 | 109.011 | 9.011 | 25.6 | 1.5699 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23166A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1730211 | 2023-06-28 |  | 14.8148 | 13.2 | 112.2333 | 12.2333 | 1.25 | 1.2918 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23179A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1650066 | 2023-07-05 |  | 263.0188 | 266 | 98.8792 | -1.1208 | 13.3 | -0.22415 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23186A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1940234 | 2023-08-01 |  | 175.4552 | 167 | 105.063 | 5.063 | 8.35 | 1.0126 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS23213A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1990166 | 2023-08-01 |  | 198.6068 | 193 | 102.9051 | 2.9051 | 10.4 | 0.53912 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS23213A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1870310 | 2023-08-08 |  | 193.8203 | 193 | 100.425 | 0.42503 | 10.4 | 0.078875 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS23219A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 1840086 | 2023-08-09 |  | 253.0862 | 266 | 95.1452 | -4.8548 | 13.3 | -0.97096 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23219A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2060142 | 2023-08-11 |  | 174.7068 | 167 | 104.6149 | 4.6149 | 8.35 | 0.92297 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS23221B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2200127 | 2023-09-01 |  | 11.0245 | 10.2 | 108.0833 | 8.0833 | 1.36 | 0.60625 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23241B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2490138 | 2023-09-26 |  | 264.3633 | 266 | 99.3847 | -0.6153 | 13.3 | -0.12306 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23269A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2350096 | 2023-09-27 |  | 10.5883 | 10.2 | 103.8069 | 3.8069 | 1.36 | 0.28551 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23269A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2280173 | 2023-10-04 |  | 464.1675 | 446 | 104.0734 | 4.0734 | 25.6 | 0.70967 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23276A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2620178 | 2023-10-04 |  | 261.8445 | 266 | 98.4378 | -1.5622 | 13.3 | -0.31244 | 2 | 3309 | Sample was diluted |  |  | µg/L | T237 | ICPMS23277A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2560088 | 2023-10-21 |  | 13.6038 | 13.2 | 103.0591 | 3.0591 | 1.25 | 0.32304 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23292A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2690094 | 2023-10-21 |  | 13.4702 | 13.2 | 102.047 | 2.047 | 1.25 | 0.21616 | 2 | 3309 |  |  |  | µg/L | T239 | ICPMS23292A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2023 | 2440134 | 2023-11-08 |  | 450.5338 | 446 | 101.0165 | 1.0165 | 25.6 | 0.1771 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS23311B
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3190183 | 2023-11-29 |  | 9.7787 | 10.2 | 95.8696 | -4.1304 | 1.36 | -0.30978 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23332A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3340080 | 2023-12-21 |  | 10.5587 | 10.2 | 103.5167 | 3.5167 | 1.36 | 0.26375 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS23354A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3390052 | 2023-12-21 |  | 10.5797 | 8.2 | 129.0207 | 29.0207 | 1 | 2.3797 | 2 | 3309 |  |  |  | µg/L | T251 | ICPMS23354A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3530158 | 2024-01-18 |  | 9.2513 | 8.2 | 112.8207 | 12.8207 | 1 | 1.0513 | 2 | 3309 |  |  |  | µg/L | T251 | ICPMS24010A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 2850090 | 2024-01-19 |  | 195.1318 | 193 | 101.1046 | 1.1046 | 10.4 | 0.20498 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS24010A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 2980081 | 2024-01-19 |  | 199.9224 | 193 | 103.5867 | 3.5867 | 10.4 | 0.66562 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS24010A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3120058 | 2024-01-19 |  | 474.6114 | 446 | 106.4151 | 6.4151 | 25.6 | 1.1176 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS24010A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3250150 | 2024-01-19 |  | 489.9617 | 446 | 109.8569 | 9.8569 | 25.6 | 1.7173 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS24010A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 2900161 | 2024-01-30 |  | 163.9741 | 167 | 98.1881 | -1.8119 | 8.35 | -0.36238 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS24025A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 2760034 | 2024-02-22 |  | 167.1328 | 167 | 100.0795 | 0.079521 | 8.35 | 0.015904 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS24051A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0370147 | 2024-02-23 |  | 11.127 | 10.2 | 109.0882 | 9.0882 | 1.36 | 0.68162 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS24053A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0230080 | 2024-03-07 |  | 167.0442 | 167 | 100.0265 | 0.026467 | 8.35 | 0.0052934 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS24065A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3460243 | 2024-03-07 |  | 263.8421 | 261 | 101.0889 | 1.0889 | 19.6 | 0.14501 | 2 | 3309 | Sample was diluted |  |  | µg/L | T249 | ICPMS24065A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 3610063 | 2024-03-07 |  | 269.9273 | 261 | 103.4204 | 3.4204 | 19.6 | 0.45547 | 2 | 3309 | Sample was diluted |  |  | µg/L | T249 | ICPMS24065A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0090034 | 2024-03-11 |  | 166.5273 | 167 | 99.7169 | -0.28305 | 8.35 | -0.056611 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS24071A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0030060 | 2024-03-11 |  | 188.7653 | 193 | 97.8059 | -2.1941 | 10.4 | -0.40718 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS24071A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0170010 | 2024-03-11 |  | 190.8003 | 193 | 98.8603 | -1.1397 | 10.4 | -0.21151 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS24071A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0440078 | 2024-03-12 |  | 465.9747 | 446 | 104.4786 | 4.4786 | 25.6 | 0.78026 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS24071A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0520079 | 2024-03-12 |  | 10.36 | 10.2 | 101.5686 | 1.5686 | 1.36 | 0.11765 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS24071A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0580158 | 2024-04-09 |  | 445.3131 | 446 | 99.846 | -0.15401 | 25.6 | -0.026832 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS24099A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0660042 | 2024-04-09 |  | 270.1178 | 261 | 103.4934 | 3.4934 | 19.6 | 0.46519 | 2 | 3309 | Sample was diluted |  |  | µg/L | T249 | ICPMS24099A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0730120 | 2024-04-24 |  | 9.6767 | 8.2 | 118.0085 | 18.0085 | 1 | 1.4767 | 2 | 3309 |  |  |  | µg/L | T251 | ICPMS24113A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0880015 | 2024-04-24 |  | 11.1896 | 8.2 | 136.4585 | 36.4585 | 1 | 2.9896 | 2 | 3309 |  |  |  | µg/L | T251 | ICPMS24113A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0800002 | 2024-04-30 |  | 271.9746 | 261 | 104.2048 | 4.2048 | 19.6 | 0.55993 | 2 | 3309 | Sample was diluted |  |  | µg/L | T249 | ICPMS24121A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 0930091 | 2024-05-08 |  | 169.3683 | 167 | 101.4181 | 1.4181 | 8.35 | 0.28363 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS24129A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1070224 | 2024-05-08 |  | 163.1711 | 167 | 97.7072 | -2.2928 | 8.35 | -0.45855 | 2 | 3309 | Sample was diluted |  |  | µg/L | T245 | ICPMS24129A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1280107 | 2024-06-27 |  | 500.2105 | 446 | 112.1548 | 12.1548 | 25.6 | 2.1176 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS24177A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1560043 | 2024-06-28 |  | 8.3558 | 8.2 | 101.9 | 1.9 | 1 | 0.1558 | 2 | 3309 |  |  |  | µg/L | T251 | ICPMS24180A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1000107 | 2024-07-09 |  | 198.3993 | 193 | 102.7976 | 2.7976 | 10.4 | 0.51916 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS24190A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1150013 | 2024-07-09 |  | 197.8691 | 193 | 102.5228 | 2.5228 | 10.4 | 0.46818 | 2 | 3309 | Sample was diluted |  |  | µg/L | T241 | ICPMS24190A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1500144 | 2024-07-10 |  | 10.3437 | 10.2 | 101.4088 | 1.4088 | 1.36 | 0.10566 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS24190A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1370189 | 2024-07-16 |  | 12.5878 | 10.2 | 123.4098 | 23.4098 | 1.36 | 1.7557 | 2 | 3309 |  |  |  | µg/L | T247 | ICPMS24192A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1420090 | 2024-07-16 |  | 513.6339 | 446 | 115.1646 | 15.1646 | 25.6 | 2.6419 | 2 | 3309 | Sample was diluted |  |  | µg/L | T243 | ICPMS24192A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1720052 | 2024-07-18 |  | 8.3456 | 8.2 | 101.7756 | 1.7756 | 1 | 0.1456 | 2 | 3309 |  |  |  | µg/L | T251 | ICPMS24200A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 1770113 | 2024-08-22 |  | 297.7253 | 261 | 114.071 | 14.071 | 19.6 | 1.8737 | 2 | 3309 | Sample was diluted |  |  | µg/L | T249 | ICPMS24233A
Zinc, unfltd, cICP-MS, recov | 01092 | PLM77 | UNF | 2024 | 2190096 | 2024-09-21 |  | 12.424 | 10.2 | 121.8039 | 21.8039 | 1.36 | 1.6353 | 2 | 3309 | Lab control sample (LCS) recovery outside of range or criteria |  |  | µg/L | T247 | ICPMS24264C